## Articles
AMDD Interpretation Matrix |  |  | 2015-10-01
|  |  | Interpretative Notes (As discussed in Dec 2015) | Points to note from 1st AMDC meeting
Article 1 General Provisions
1 |  | Each Member State shall undertake all necessary measures to ensure that only medical devices which conform to the provisions of this ASEAN Agreement on Medical Device Directive (hereinafter referred to as “Agreement”) and its Annexes may be placed on the markets of that Member State. | Medical Devices conforming to the AMDD and its annexes shall only be given marketing authorisation in the AMS.
2 |  | Subject to the provisions of this Agreement each Member State shall require that the person responsible for placing the medical device in that Member State or the authorised representative shall register the medical device with the Regulatory Authority of that Member State. | The person responsible for placing the MD on the market in the MS or authorised representative (designated by the product owner of the MD) shall be required to register the medical devices with the RA of the MS.
3 |  | Subject to the provisions of this Agreement, each Member State shall require that the person responsible for placing the medical device in that Member State or the authorised representative shall be licensed by the Regulatory Authority of that Member State before placing the medical device in that Member State. | The person responsible for placing the MD on the market in the MS or the authorised representative (designated by the product owner of the MD) shall be licensed by the RA of the MS.
Article 2 Definitions and Scope
1 |  | This Agreement shall apply to medical devices and their accessories. For the purposes of this Agreement, accessories shall be treated as medical devices in their own right. Both medical devices and accessories shall hereinafter be referred to as devices. | The scope of AMDD includes both the medical devices and their accessories.
| a | “medical device” shall mean any instrument, apparatus, implement, machine, appliance, implant, in vitro reagent and calibrator, software, material or other similar or related article: (i) intended by the product owner to be used, alone or in combination, for human beings for one or more of the specific purpose(s) of: (A) diagnosis, prevention, monitoring, treatment or alleviation of disease; (B) diagnosis, monitoring, treatment, alleviation of or compensation for an injury; (C) investigation, replacement, modification, or support of the anatomy or of a physiological process; (D) supporting or sustaining life; (E) control of conception; (F) disinfection of medical devices; and (G) providing information for medical or diagnostic purposes by means of in vitro examination of specimens derived from the human body; (ii) which does not achieve its primary intended action in or on the human body by pharmacological, immunological or metabolic means, but which may be assisted in its intended function by such means. | Definition
| b | “accessory” means an article that is intended specifically by its product owner to be used together with a particular medical device to enable or assist that device to be used in accordance with its intended purpose. | Definition
| c | “adverse event” means either a malfunction or a deterioration in the characteristics or performance of a supplied medical device or use error, which either has caused or could have caused or contributed to death, or injury to health of patients or other persons. | Definition
| d | “authorised representative” means any person in a Member State who, explicitly designated by the product owner, acts and may be addressed by authorities and bodies in a Member State instead of the product owner with regard to the latter’s obligations under this Agreement, and relevant laws and regulations of the Member State. | Definition
| e | “authorised distributor”, in relation to the placing on the market of a medical device, means any person who has been authorised by the product owner or authorised representative to distribute the medical device in that Member State. | Definition
| f | “custom-made medical device” means any device specifically made in accordance with a duly qualified medical practitioner's written prescription which gives, under his responsibility, specific design characteristics and is intended for the sole use of a particular patient. For the purposes of this definition, a duly qualified medical practitioner is defined as a person who is duly qualified by the relevant laws and regulations of the Member State where the custom-made medical device is used. | Definition
| g | “device intended for clinical investigation” means any device intended for use by a duly qualified medical practitioner when conducting clinical investigations as referred to in Annex 8 (Clinical Investigation), in an adequate human clinical environment. For the purposes of conducting of clinical investigation, a duly qualified medical practitioner is defined as a person who is duly qualified by the relevant laws and regulations of the Member State where the clinical investigation is carried out, and by virtue of his professional qualifications, is authorised to carry out such investigation. | Definition
| h | “Field Safety Corrective Action (FSCA)” means any action taken by a product owner to reduce a risk of death or serious deterioration in the state of health associated with the use of a medical device. This may include: (i) the return of a medical device to the product owner or its representative; (ii) device modification which may include: (a) retrofit in accordance with the product owner's modification or design change; (b) permanent or temporary changes to the labelling or instructions for use; (c) software upgrades including those carried out by remote access; (d) modification to the clinical management of patients to address a risk of serious injury or death related specifically to the characteristics of the device. (iii) device exchange; (iv) device destruction; (v) advice given by product owner regarding the use of the device. | Definition
| i | “intended purpose” means the use for which the medical device is intended according to the specifications of its product owner as stated on any or all of the following: (i) the label of the medical device; (ii) the instructions for use of the medical device; (iii) the promotional materials in relation to the medical device. | Definition
| j | “in vitro diagnostic (IVD) medical device” means any reagent, reagent product, calibrator, control material, kit, instrument, apparatus, equipment or system, whether used alone or in combination with any other reagent, reagent product, calibrator, control material, kit, instrument, apparatus, equipment or system, that is intended by its product owner to be used in vitro for the examination of any specimen, including any blood or tissue donation, derived from the human body, solely or principally for the purpose of providing information: (i) concerning a physiological or pathological state or a congenital abnormality; (ii) to determine the safety and compatibility of any blood or tissue donation with a potential recipient thereof; or (iii) to monitor therapeutic measures; and includes a specimen receptacle. | Definition
| k | “manufacture”, in relation to a medical device, means to make, fabricate, produce or process the medical device and includes: (i) any process carried out in the course of so making, fabricating, producing or processing the medical device; and/or (ii) the packaging and labelling of the medical device before it is supplied. | Definition
| l | “person” means a natural person or a legal entity including a corporation, a partnership or association duly established pursuant to the prevailing laws and regulations of Member States. | Definition
| m | “physical manufacturer”, in relation to a medical device, means any person who performs the activity of manufacture. | Definition
| n | “placing on the market” means the making available in return for payment or free of charge of a medical device other than a device intended for clinical investigation, with a view to distribution and/or use on the market of a Member State. | Definition
| o | “product owner”, in relation to a medical device, means any person who: (i) supplies the medical device under his own name, or under any trade mark, design, trade name or other name or mark owned or controlled by him; and (ii) is responsible for designing, manufacturing, assembling, processing, labelling, packaging, refurbishing or modifying the medical device, or for assigning to it a purpose, whether those tasks are performed by him or on his behalf. | Definition
| p | “putting into service” means the stage at which a medical device has been made available to the final user as being ready for use on the market of a Member State for its intended purpose. | Definition
| q | “refurbished medical device” means a medical device of which the whole or any part thereof has been substantially rebuilt, whether or not using parts from one or more used medical devices of that same kind, so as to create a medical device that can be used for the purpose originally intended by the product owner of the original medical device, and which may have had the following work carried out on it: (i) stripping into component parts or sub-assemblies; (ii) checking their suitability for reuse; (iii) replacement of components/sub-assemblies not suitable for reuse; (iv) assembly of the reclaimed and/or replacement components/sub-assemblies; (v) testing of the assembled device against either original or revised release criteria; or (vi) identifying an assembled medical device as a refurbished medical device. | Definition
| r | register” means to obtain marketing approval for a medical device from the Regulatory Authority of a Member State in order to place the medical device on the market of that Member State. | Definition
| s | ”Regulatory Authority” means the regulatory authority or entity of that Member State which exercises a legal right to control the import, manufacture, export, distribution, transfer, use and sale of medical devices within that Member State’s jurisdiction and which may take regulatory action to ensure that the products marketed within its jurisdiction comply with regulatory requirements. | Definition
| t | “sponsor” means an individual or organisation taking responsibility and liability for the initiation or implementation of a clinical investigation. | Definition
2 |  | This Agreement shall not apply to the following:
| a | human blood, plasma or blood cells of human origin or to medical devices which incorporate at the time of placing on the markets of Member States such human blood, plasma or blood cells of human origin, except if: (i) it is incorporated in an IVD medical device, or (ii) it is incorporated in a medical device as a human blood derivative with an action ancillary to that of the medical device. | Products containing blood and blood cells of human origin are not subject to controls as MD. Exceptions are IVDs and combination MDs where the human origin material has an ancillary role to the MD intended purpose.
| b | transplants or tissues or cells of human origin nor to products incorporating or derived from tissues or cells of human origin, except if it is incorporated in an IVD medical device; or | Products containing cells or tissues of human origin are not subject to controls as MD. Exceptions are IVDs and combination MDs where the human origin material has an ancillary role to the MD intended purpose.
| c | transplants or tissues or cells of animal origin, unless: (i) it is incorporated in an IVD medical device, or (ii) it is a medical device manufactured utilizing animal tissue which is rendered non-viable or non-viable products derived from animal tissues or cells. “Non-viable” means in relation to a biological entity, an entity that is incapable of growth, development and reproduction. | Products containing "viable" or "live" cells or tissues of animal origin are not subject to controls as MD. Exceptions are IVDs and non-viable animal tissues/cells.
Article 3 Essential Principles of Safety and Performance of Medical Device
|  | Medical devices shall meet the essential principles set out in Annex 1 (Essential Principles of Safety and Performance of Medical Devices), as may be applicable, taking account of the intended purpose of the medical device concerned. | All MDS are required to conform to the essential requirements for safety and performance of MDs - Annex 1 of AMDD
Article 4 Classification of Medical Devices
1 |  | Medical devices shall be classified into the following four classes, in accordance with risk classification rules set out in Annex 2 (Risk Classification Rules for Medical Devices other than IVD Devices) and Annex 3 (Risk Classification Rules for IVD Devices): Class Risk Level A Low risk B Low-moderate risk C Moderate-high risk D High risk | All MDs shall be classified based on the risk classification rules laid out in the Annex 2 and 3 of the AMDD document into 4 risk classes i.e. A,B,C and D. Class A bieng the lowest risk class and Class D the highest.
2 |  | In the event that a medical device may be assigned into two or more classes of medical devices, the Regulatory Authority of the Member State shall assign the medical device into such of those classes as represents the highest health risk posed to an end-user of the medical device. | MDs that may be assigned to two or more risk classes based on the rules described in Annex 2 and 3, the AMS shall assign the MD the highest of the risk class possible based on the applicable rules to the MD.
3 |  | In the event that a medical device is designed to be used in combination with another medical device, each of the medical devices shall be classified separately. | Where multiple MDs are designed to be used in combination, the risk classification shall be assiged individually to each of these MDs using the rules in Annex 2 and 3.
4 |  | In the event the medical device has two or more intended purposes, the medical device shall, subject to paragraph 3 of this Article, be assigned into a class of medical devices having regard to the most critical intended purpose of the medical device. | Where a MD has multiple intended purposes, the MD shall be classifed considering the shall the most critical intended purpose.
5 |  | In the event of a dispute between a Member State and any person in the classification of a medical device, the Regulatory Authority of that Member State shall decide on the proper classification of the medical device concerned, whose decision shall be final. | Where there's a disagreement in MD risk classification between an MS and any person (natural person/legal entity), the final risk classificaiton of the MD shall be based on the decision of regulatory authority of the MS.
|  | Member State that reclassifies or differs in its application of the classification rules set out in Annex 2 (Risk Classification Rules for Medical Devices other than IVD Devices) and Annex 3 (Risk Classification Rules for IVD Devices) shall notify, with the reasons thereof, to the ASEAN Medical Device Committee (hereinafter referred to as “AMDC”) to be established pursuant to Article 14 of this Agreement, of such measures taken. | Under circumstances where the AMS differs in its application of the classification rule set out in the Annex 2 and 3 of the AMDD, the member state shall notify with appropriate reasons to the AMDC.
Article 5 Conformity Assessment of Medical Devices
1 |  | A medical device, which is required to be assessed by a Member State and intended to be placed on the market of that Member State, shall be assessed by the Regulatory Authority of that Member State, or any appointed bodies recognised by that Member State, as the case may be, for conformity and compliance with at least the requirements laid down in this Agreement unless the medical device has been exempted from the requirement for registration under paragraph 2 of Article 6. | MDs that are required to be reviewed shall be assessed by the AMS or any appointed bodies (e.g. CAB) that the MD is in compliance with the requirements of the AMDD (e.g. Essential Requirements, technical documentation etc.) prior to being placed on the market of the AMSs. Exception to this may be devices under circumstances described in Article 6 (i.e. custom-made MDs or MDs required in the interest of public health)
2 |  | Member States shall put in place an appropriate system for the conformity assessment of medical devices as referred to paragraph 1 of this Article. | An appropriate system for the conformity assessment of MDs shall be put in place in the AMS.
Article 6 Registration and Placement on the Market
1 |  | A medical device, which is required to be assessed by a Member State and has been assessed by the Regulatory Authority of that Member State or any appointed bodies recognised by that Member State to be in conformity and in compliance with the requirements laid down in this Agreement may be placed on the market of that Member State. | MDs that are required to be reviewed shall be assessed by the member state or any appointed bodies (e.g. CAB) that the MD is in compliance with the requirements of the AMDD (e.g. Essential Requirements, technical documentation etc.) prior to being placed on the market of the AMSs.
2 |  | A medical device to be placed on the market of a Member State shall be registered with the Regulatory Authority of that Member State. The Regulatory Authority of the Member State may exempt certain medical devices from the requirement for registration where appropriate. | MDs to be placed on market of the AMS shall be registered with the RA of the member state. However, the AMS may choose to exempt certain MDs from this requirement (e.g. low risk MDs)
3 |  | Member States shall put in place an appropriate system for the registration of medical devices with the Regulatory Authority of that Member State. | An appropriate system for registration of MDs shall be put in place in the AMSs.
4 |  | Custom-made medical devices shall not be subjected to product registration requirements. | Custom-made medical devices have an exception from product registration requirement.
5 |  | Notwithstanding paragraphs 1 and 2 of this Article, the Regulatory Authorities of a Member State may, pursuant to a duly justified request or on its own initiative, authorise the use within the territory of that Member State, of medical devices which have not undergone registration with the Regulatory Authority and where such use is in the interest or protection of public health. | Under circumstances where the RA determines that in the interest of public health (e.g. urgent clinical need for a patient(s)) there is a need for the use of an unregistered MD, the RA may proceed to allow for such authorisation.
Article 7 Licensing of Person Responsible for Placing Medical Devices on the markets of Member States
|  | Each Member State shall require a person who is responsible for placing medical devices on the market to be licensed by the Regulatory Authority of that Member State before the medical devices are placed on the market of that Member State. Member States shall put in place an appropriate system for the licensing of persons responsible for placing medical devices on their markets. | The person (a natural person/a legal entity) responsible for placing the MD on the market shall be licensed by the RA of the MS. An appropriate system for licensing of a person placing a MD on the market shall be put in place in the MS.
Article 8 Technical Documents for Medical Devices
|  | Member States shall undertake appropriate measures to adopt and implement the following common technical documents: (a) Annex 4 - ASEAN Common Submission Dossier Template (CSDT); (b) Annex 5 - Post Marketing Alerts System (PMAS) Requirements; and (c) Annex 6 - Harmonized set of elements for a Product Owner’s or Physical Manufacturer’s Declaration of Conformity (DoC). | Annex 4 adoption (including template) is mandatory. Applies to all MDs, except for IVD, MS may choose to adopt this CSDT or prescribe another format. Ref to Annex 5 sec 1.2, which states "MS may adopt the recommended post-market alerting system requirements in this Annex or prescribe their own post-market alerting system requirements." - Annex 5 PMAS is optional Annex 6 adoption (including DoC template) is mandatory.
Article 9 Reference to Technical Standards
1 |  | Medical devices which conform to either the relevant technical standards recognised by the AMDC or other technical standards accepted by the Regulatory Authority of a Member State for the medical device to be placed in the market of that Member State shall be deemed to comply with the applicable essential principles referred to in Article 3. | MD that complies with standards recognised by AMDC/MS would be considered to be in compliance to EP. [Note that compliance to standards is voluntary. However, need to take note of the list of standards and the respective version to be adopted by AMDC (and how AMDC deals with subsequent revisionsof standards)]
2 |  | The AMDC may revise by consensus, the recognised technical standards referred to in paragraph 1 of this Article. | Any changes to the list of standards would required consensus from all MS.
Article 10 Labelling
1 |  | A medical device shall be labelled in accordance with the requirements of the Member State prior to placing on the market in that Member State. | MS to decide on labelling requirement.
2 |  | Member States may set the labelling requirements for a medical device in accordance with Annex 7 (Labelling Requirements) or as deemed appropriate by the Member States. | Adoption of labelling requirement in Annex 7 is optional.
3 |  | Member States may set the requirement for having the label of a medical device in their national languages. | MS have the rights to mandate MDs to be labelled based on individual national languages.
Article 11 Medical Device Claims
1 |  | Medical device claims shall be subject to regulatory control of Member States. | MD claims for marketing the device in the MS will be subject to individual MS clearance in accordance to national legislation.
2 |  | As a general rule, claimed benefits of a medical device shall be justified by substantial evidence and/or by the medical device composition/formulation/component or preparation itself in accordance with the requirements as set out in Annex 1 (Essential Principles for Safety and Performance of Medical Devices). | Claims shall be substantiated by evidences and shall demonstrate compliance with the Annex 1 of AMDD i.e. EP
Article 12 Post-marketing Alert System
1 |  | Member States shall take the necessary steps to ensure that any information brought to their knowledge, in accordance with the provisions of this Agreement, regarding the incidents involving a medical device as mentioned below is recorded and evaluated when appropriate: | MS to ensure incident involving a MD that meets (a) or (b) condition is recorded and evaluated "when appropriate"
| a | any malfunction or deterioration in the characteristics or performance of a medical device, as well as any inadequacy in the labelling or the instructions for use which might lead to or might have led to the death of a patient or user or to a serious deterioration in his state of health; | a. Issue relates to malfunction/performance/labelling, that which might lead to death/serious injury.
| b | any technical or medical reason in relation to the characteristics or performance of a medical device for the reasons referred to in subparagraph (a), leading to product recall of medical devices of the same type by the product owner, authorised representative, authorised distributor or person responsible for placing medical device into the market. | Issue in (a) lead to recall of MD of "same type". [Does MD of "same type" refers to normal FSCA resulted from AE quoted in (a)?]
2 |  | After carrying out an assessment, if possible together with the product owner, a Member State shall inform the other Member States of the incidents referred to in paragraph 1 of this Article for which relevant measures have been taken or are contemplated. | After assessment, MS to inform other MS on measures taken or are contemplated. [Does this imply that it is mandatory to inform other MS of all AE of death/serious injury & FSCA resulted from thereof?] | MS to discuss further on the requirements - e.g. type of incidents, mechanism, scope
3 |  | Each Member State shall require that any person, who is responsible for the manufacture or placing the medical devices on the market of that Member State, to: | MS is require to mandate any person (manufacturer/ placing on market) to perform the (a) - (d)
| a | keep all relevant records pertaining to the traceability of the medical device, for such period and format as the Regulatory Authority in the Member State may stipulate; | a. record keeping for traceability purpose Period, format, and info not prescribed.
| b | produce such records for inspection when required by the Regulatory Authority in the Member State; | b. produce records in a. when requested
| c | inform the Regulatory Authority, within the stated prescribed time and format of the Regulatory Authority in the Member State, where he becomes aware of any adverse event that has arisen or can arise from the use of the medical device placed on the market in the Member State; and | c. Inform authority of AE (define in Article 2) Time and format is not prescribed. Refer to Article 8 comment, Annex 5 is optional.
| d | inform the Regulatory Authority, within the stated prescribed time and format of the Regulatory Authority in the Member State, when he performs or intends to perform FSCA on a medical device placed on the market in the Member State. | d. Inform authority of any FSCA (define in Article 2) Time and format is not prescribed. Refer to Article 8 comment, Annex 5 is optional.
Article 13 Clinical Investigation
|  | Member States shall put in place an appropriate system for the conduct of clinical investigation of medical devices, taking into account the Helsinki Declaration adopted by the 18th World Medical Assembly in Helsinki, Finland, in 1964, and any subsequent amendments or revisions to this Declaration by the World Medical Association. It is acknowledged that all measures relating to the protection of human subjects are required to be carried out in accordance with the spirit of the Helsinki Declaration. This includes every step in the clinical investigation from first consideration of the need and justification of the study to publication of the results, which may include the following requirements: | MS shall put in place an "appropriate system" for conduct of clinical trial for MD and take into account Helsinki declaration. AMDD does not prescribe what system needs to be in place with the condition that human subjects are protected with the spirit of the Helsinki Declaration.
1 |  | In the case of medical devices intended for clinical investigation, the Regulatory Authority of the Member State may require the product owner, or his authorised representative, or the sponsor of the clinical investigation in a Member State, as the case may be, to follow the procedure referred to in Annex 8 (Clinical Investigation) and register with the Regulatory Authority of that Member State in which the investigations are to be conducted. | MS could request the appropriate party to conduct clinical trial per Annex 8 and register with the MS where trial is to be conducted. Adoption of Annex 8 is optional.
2 |  | The Regulatory Authority of the Member State may require that the clinical investigations be conducted in accordance with the provisions of Annex 8 (Clinical Investigation). | Adoption of Annex 8 is optional.
3 |  | The Regulatory Authority of that Member State may require the product owner or his authorised representative, or the sponsor of the clinical investigation in a Member State, as the case may be, to submit or make available on request, as deemed appropriate, the report referred to in Annex 8 (Clinical Investigation). | MS "may" request for a report of the trial, as referred in Annex 8.
4 |  | Where a clinical investigation is refused or halted by a Member State, that Member State may communicate its decision and the grounds thereof to all Member States and the AMDC. Where a Member State has called for a significant modification or temporary interruption of a clinical investigation, that Member State may inform all Member States and the AMDC concerned about its actions and the grounds for the actions taken. | Where a clinical investigation is rejected/ modified significantly/ cease by a MS, the MS "may" choose to inform other MS of its decision and reason to all MS and AMDC.
5 |  | The Regulatory Authority of a Member State may require the product owner or his authorised representative, or the sponsor of the clinical investigation in a Member State, as the case may be, to notify of the end of the clinical investigation, with justification(s) in case of temporary suspension or of early termination. In the case of early termination of the clinical investigation on safety grounds, this notification may be communicated to the Regulatory Authority of all Member States where the clinical investigation is carried out. | MS could request the appropriate party to notify the authority in case of temporary suspension/ early termination, with justifications. For termination due to safety issue, info "may"be communicated to other MS where trial is conducted.
Article 14 Institutional Arrangements
1 |  | The AMDC shall be established with the overall responsibility of coordinating, reviewing and monitoring the implementation of this Agreement and shall comprise representatives from the Regulatory Authority of each Member State. | Establishment of AMDC and its role. Self-explanatory.
2 |  | The ASEAN Consultative Committee for Standards and Quality (ACCSQ) and the ASEAN Secretariat shall provide support in coordinating and monitoring the implementation of this Agreement and assist the AMDC in all matters relating thereto. | Define ACCSQ and ASEAN Sec supporting role to AMDD implementation and in assisting AMDC in this regard.
3 |  | The AMDC may establish an ASEAN Medical Device Technical Committee (AMDTC) to assist the AMDC in reviewing the technical and safety issues. | AMDC can choose to establish AMDTC, to review technical related matters in relation to AMDD.
Article 15 Safeguard Clauses
1 |  | Member States acknowledge that a medical device placed on the market of Member States shall not compromise the health or safety of patients, users or, where applicable, other persons, when applied under normal or reasonably foreseeable conditions of use, taking account, in particular, of the medical device’s presentation, packaging, its labelling, instructions for its use and where appropriate, disposal, warning statements as well as any other indication or information provided by the product owner or his authorised representative or by any other person responsible for placing the medical device on the market. | MS to acknowledge that MD placed on the market of MS shall not compromise the health/safety to any person, based on proper use in accordance to PO's instruction. Excludes off label use.
2 |  | Where a Regulatory Authority ascertains that a medical device placed on the market of a Member State, when correctly installed, maintained and used for its intended purpose, may compromise the health or safety of patients, users or, where applicable, other persons, it shall take all appropriate interim measures to withdraw such medical device from the market or prohibit or restrict their being placed on the market or put into service. That Member State shall immediately inform the other Member States of any such measures, indicating the reasons for its decision and, in particular, whether this is due to: | Provide MS the rights to withdraw the use of a MD from market due to safety and/or performance issue and responsibility to inform other MS of the event especially if the reason is due to points a,b or c below. | To be discussed in more details at the next HOD Meeting when reviewing the draft PMS Guidelines. For further discussion - on sharing requirements/ mechanism.
| a | the medical device failing to meet the essential principles set out in Annex 1 (Essential Principles for Safety and Performance of Medical Devices); | a. Non compliance with EP
| b | incorrect application of the technical standards referred to in Article 9, in so far as it is claimed that the technical standards have been applied; or | b. Incorrect claim of complaince to technical standards
| c | shortcomings in the technical standards themselves. | c. limitations in the standards complied with
Article 16 Confidentiality
|  | Without prejudice to the existing national laws and regulations, Member States shall require that all the parties involved in the application of this Agreement are bound to observe confidentiality with regard to all confidential information obtained in carrying out their tasks. This confidentiality obligation does not however affect or prevent: | In addition to individual National laws, MS shall ensure confidentiality of information obtained while implementing AMDD requirements. Note: information may be obtained (in carrying out task) includes pre-market and post-market evaluations and investigations in the MS. Under this Article, this confidentiality obligation will still allow:
| a | the disclosure of any information by any Member State to another Member State pursuant to any provision in this Agreement; | a. any MS to disclose/share necessary information to another Member State, in accordance to any of AMDD requirements. Does this refer to provisions for sharing information between MSs under Articles 12, 13 or 15 for example? | Article 12 to be discussed with post-market guidelines in next meeting. Article 13 - sharing is optional.
| b | the disclosure of any confidential information with the permission of the person from whom the information was obtained; | b. MS may share confidential information with permission from the source of the information.
| c | the disclosure of information for the purposes of the administration or enforcement of the requirements of this Agreement including the disclosure of any information to any adviser engaged by the Regulatory Authority of a Member State to advise on any aspect of the medical device to which the information relates and the dissemination of warnings or information for the public interest; | MS may share information on any medical device where necessary for regulatory action (e.g. enforcement, to inform public, etc), including MS's adviser.
| d | the disclosure of information for the purposes of assisting in any investigation or prosecution of any offence under the national law of a Member State; and | MS may share information on any medical device where necessary for regulatory action, includes investigation or prosecution; in other MSs.
| e | any requirement by any court or the provisions of any national law of a Member State to provide information. | MS may share information when required by Court or any National law of MSs.
Is there a need for a formal procedure to be established for above sharing information?
Article 17 Special Cases
1 |  | A Member State may refuse to register or prohibit the marketing of a medical device in its market or subject it to special conditions or different controls, as it deems appropriate, although the medical device complies with the requirements of the Agreement, for reasons specific to religious or cultural sensitivity. | On religious or cultural grounds, MSs exercise discretion to apply special conditions or controls to a medical device, or even refuse to register or prohibit the device.
2 |  | A Member State may refuse or prohibit a refurbished medical device to be placed on its market or put into service, as it deems appropriate, even if such medical device complies with the requirements of the Agreement. | For refurbished medical devices, MSs can exercise discretion to refuse to register or prohibit such products although they may be in compliance with the requirements under AMDD.
3 |  | Nothing in this Agreement shall be construed to limit the authority of a Member State to determine, through its legislative, regulatory and administrative measures, the level of protection it considers appropriate for safety; for protection of human, animal, or plant life or health; for the environment and for consumers. | MSs can exercise discretion to impose additional requirements and controls (National laws) beyond this AMDD on certain devices in the interest of protection of public/ animal/plant/environmental good.
4 |  | Nothing in this Agreement shall be construed to limit the authority of a Member State to take all appropriate and immediate measures whenever it ascertains that a medical device may: (a) compromise the public health or safety in its territory; (b) not meet the legislative, regulatory, or administrative provisions within the scope of this Agreement; or (c) otherwise fail to comply with a requirement within the scope of this Agreement. | MSs can take any necessary immediate action, if device poses a threat to public health, not meet laws and provisions (including administrative) to meet the AMDD, and not comply with requiremets of Agreement.
5 |  | A Member State that places a restriction or ban on specific medical devices shall notify the other Member States with the reasons thereof, together with a copy to the AMDC of such measures taken. | MSs must notify other Member States, and also the AMDC, the reasons and explanation, when it restrics or bans a device that is in compliance with the AMDD requirements.
Article 18 Implementation
1 |  | Member States shall undertake appropriate measures to implement this Agreement. | MSs to implement requirements in this AMDD, adequately and systematically.
2 |  | Member States shall undertake appropriate measures to ensure that the technical infrastructures necessary are in place to implement this Agreement. | MSsto provide adequate infrastructure to implement requirements of this AMDD. This would likely include training, capacity building for RA and stakeholders, systems and tools (e.g. for registration, validation)
3 |  | Member States shall ensure that the texts of such provisions of national laws, which they adopt in the field governed by this Agreement are communicated to the other Member States with a copy to the ASEAN Secretariat, that shall promptly notify the AMDC. | MSs to share copy of national laws for meeting the AMDD.
4 |  | Member States shall ensure that post marketing surveillance is in place and shall have full authority to enforce the law on medical devices found to be not complying with this Agreement. | Member States implement post-market surveillance. Ensure adequate authority to enforce requirements in AMDD and enforce on MDs noncompliant to these requirements.
Article 19 Revisions, Modifications and Amendments
1 |  | The provisions of all or any part of this Agreement may be revised, modified or amended, by written agreement of all Member States. | With written agreement of all MSs, AMDD elements may be edited.
2 |  | Notwithstanding paragraph 1 of this Article, the Annexes of this Agreement may be revised, modified or amended subject to the endorsement of AMDC. | Annexes of AMDD, may be edited subject to the endorsement of AMDC only.
3 |  | Any revision, modification or amendment shall not prejudice the rights and obligations arising from or based on this Agreement prior and up to the date of such revision, modification or amendment. | Edits will not affect retrospect positions/policies.
Article 20 Dispute Settlement
|  | The ASEAN Protocol on Enhanced Dispute Settlement Mechanism signed on 29 November 2004 in Vientiane, Lao PDR and amendments thereto, shall apply to the settlement of disputes concerning the interpretation or implementation of this Agreement. | Settle disputes relating to the interpretation and implementation of the AMDD as defined by the ASEAN Protocol.
Article 21 Reservations
|  | Member States shall make no reservation with respect to any of the provisions of this Agreement. | MSs to implement AMDD in totality.
Article 22 Entry into Force
1 |  | This Agreement shall be subject to ratification and/or acceptance by Member States in accordance with their internal domestic requirements. | AMDD shall be ratified and accepted/implemented by the MS based on their domestic requirments.
2 |  | This Agreement shall enter into force on 1 January 2015 and shall be in force only among the Member States that have ratified and/or accepted it. | This Agreement shall enter into force from 1 Jan 2015, and shall be in force only among MSs that have ratified and/or accepted it.
3 |  | The Secretary-General of ASEAN shall promptly notify all Member States of the notifications or deposit of each instrument of ratification and/or acceptance referred to in paragraph 1 of this Article. | Duty of Secretary-General of ASEAN to notify all MSs of the information regarding the ratification or acceptance of AMDD in other MS.
Article 23 Annexes
|  | The Annexes to this Agreement constitute an integral part of this Agreement. | Annexes are part of the Agreement
Article 24 Depositary
|  | This Agreement shall be deposited with the Secretary-General of ASEAN, who shall promptly furnish each Member State a certified copy thereof. | This Agreement shall be deposited with the Secretary-General of ASEAN, and provide each Member State with a certified copy.

## Annex 1 EP
ARTICLE 3 ESSENTIAL PRINCIPLES OF SAFETY AND PERFORMANCE OF MEDICAL DEVICE
Medical devices shall meet the essential principles set out in Annex 1 (Essential Principles of Safety and Performance of Medical Devices), as may be applicable, taking account of the intended purpose of the medical device concerned
Conformity Reference - Examples (For discussion)
General Requirements
1 |  | Medical devices shall be designed and manufactured in such a way that, when used under the conditions and for the purposes intended and, where applicable, by virtue of the technical knowledge, experience, education or training of intended users, they will not compromise the clinical condition or the safety of patients, or the safety and health of users or, where applicable, other persons, provided that any risks which may be associated with the use of the medical device for its intended purpose constitute acceptable risks when weighed against the intended benefits to the patient and are compatible with a high level of protection of health and safety. | Standards: ISO 13485 (QMS), ISO 14971 (Risk Management), ISO 10993 (Biocompatibility) Others: In-house validation reports
2 |  | The solutions adopted by the product owner for the design and manufacture of the medical devices shall conform to safety principles, taking account of the generally acknowledged state of the art. In selecting an appropriate solution for the design and manufacture of a medical device so as to minimise any risks associated with the use of the medical device , the product owner shall apply the following principles: • identify any hazard and associated risk arising from the use of the medical device for its intended purpose, and any foreseeable misuse of the medical device, • eliminate or reduce risks as far as reasonably practicable through inherently safe design and manufacture, • if appropriate, ensure that adequate protective measures are taken, including alarms if necessary, in relation to any risk that cannot be eliminated, and • inform users of any residual risks. | Standards: ISO 14971 (Risk Management), ISO 10993 (Biocompatibility), Product specific standards (where available) Others: Risk management Report: No unacceptable residual risks In-house validation reports, Device Master Records (where applicable) Where possible design related risk control measures (e.g. alarms, auto shut down mechanisms etc.) Labeling information - least preferred risk mitigation measure - to address residual acceptable risks only
3 |  | Medical devices shall achieve the performance intended by the product owner and be designed, manufactured and packaged in such a way that they are suitable for one or more of the functions within the scope of the definition of a medical device. | Standards: Product specific standards (where available) Others: Finished product specifications, Validation reports
4 |  | The characteristics and performances referred to in Clauses 1, 2 and 3 shall not be adversely affected to such a degree that the health or safety of the patient or the user and, where applicable, of other persons are compromised during the lifetime of the medical device, as indicated by the product owner, when the medical device is subjected to the stresses which can occur during normal conditions of use and has been properly maintained and calibrated, if appropriate, in accordance with the product owner’s instructions. | Standards: Product specific standards (where available) Others: Device Maser Records, Finished product specifications, Validation report - Includes shelf life and stability testing, Functional tests Appropriateness of the recommended maintenance and calibration instructions presented for the user
5 |  | The medical devices shall be designed, manufactured and packed in such a way that their characteristics and performances, when it is being used for its intended purpose, will not be adversely affected during its transport and storage, if the transport and storage is carried out in accordance with the instructions and information provided by the product owner. | Standards: Product specific standards (where available) Others: Finished product specifications, Storage Recommendations Validation Reports including stability testing - Shipping, transportation and temperature cycling protocols with appropriate functional tests on the device. Labels presenting the appropriate storage conditions and special instrucitons if any to the user
6 |  | The benefits must be determined to outweigh any undesirable side effects for the performances intended. | Standards: ISO 14971 (Risk Management) Others: Risk management Report: No unacceptable residual risks Any undersirable side effect or residual risk must be assessed and justified to constitute an acceptable risk when weighed against the device intended purpose and performance.
7 |  | Medical devices shall require clinical evidence, appropriate for the use and classification of the medical device, demonstrating that the medical device complies with the applicable provisions of the essential principles. A clinical evaluation shall be conducted. | Standards: ISO 14155 (Clinical evidence for MDs) Others: Clinical Evaluation Report inluding clinical evidence gathered from one or more of the following: Scientific Literature Review, Clinical Experience and Clinical evaluation studies.
Design and Manufacturing Requirements
8 | Chemical, Physical and Biological Properties
8.1 |  | The medical devices shall be designed and manufactured in such a way as to ensure the characteristics and performance requirements referred to in Clauses 1 to 6 of the 'General Requirements' are met. Particular attention shall be paid to: • the choice of materials used, particularly as regards toxicity and, where appropriate, flammability, • the chemical and physical properties of the material used, • the compatibility between the materials used and biological tissues, cells, body fluids, and specimens, taking account of the intended purpose of the medical device, • the choice of materials used shall reflect, where appropriate, matters such as hardness, wear and fatigue strength. | Standards:ISO 10993 (Biocompatibility), Product specific standards (where available) Others: Mechanical or bench testing, Other validation reports, Material Data Sheets
8.2 |  | The medical devices shall be designed, manufactured and packed in such a way as to minimise the risk posed by contaminants and residues to the persons involved in the transport, storage and use of the medical devices and to patients, taking account of the intended purpose of the product. In minimising risks, particular consideration shall be given to the duration and frequency of any tissue exposure associated with the transport, storage or use of the medical device. | Standards:ISO 10993 (Biocompatibility), ISO 14971 (Risk Management) Others: Packaging Validation, Risk management plan, Other validation reports
8.3 |  | The medical devices shall be designed and manufactured in such a way that they can be used safely with the materials, substances and gases with which they enter into contact during their normal use or during routine procedures; if the medical devices are intended to administer medicinal products they shall be designed and manufactured in such a way as to be compatible with the medicinal products concerned according to the provisions and restrictions governing these medicinal products and that the performance of the medicinal product is maintained in accordance with the intended purpose of the medicinal product. | Standards:ISO 14971 (Risk Management), ISO 10993 (Biocompatibility), Others: Risk management plan, Compatibility validation with medicinal products including reference to pharmacopeia or other standards for drugs where relevant, Other validation reports
8.4 |  | Where a medical device incorporates, as an integral part, a substance which, if used separately, may be considered to be a medicinal product as defined in the relevant legislation that applies and which is liable to act upon the body with action ancillary to that of the medical device, the safety, quality and performance of the medical device as a whole shall be verified, as well as the safety, quality and efficacy of the incorporated substance in relation to the intended purpose of the medical device. For the purposes of this paragraph, “medicinal product” includes any stable derivative of human blood or human plasma. | Applicable for Device-Drug combination products - devices incorporating a regstrable drug (including biologics such as human blood derivatives) in an ancilary role (e.g. drug eluting stent) Others: Validation reports including clinical evidence- required to substantiate the safety, quality and efficacy of the combined MD in addition to data to support the ancillary substance (i.e. drug or biologic)
8.5 |  | The medical devices shall be designed and manufactured in such a way as to reduce as far as reasonably practicable and appropriate the risks posed by substances that may leach or leak from the medical device. | Standards: ISO 10993 (Biocompatibility) Others: Device design and composition, Packaging validation, Material Data Sheets
8.6 |  | Medical devices shall be designed and manufactured in such a way as to reduce as far as reasonably practicable and appropriate risks posed by the unintentional ingress or egress of substances into or from the medical device taking into account the nature of the environment in which the medical device is intended to be used. | Others: Finished product Specifications and batch testing as applicable, Stability validation including shipping/transportation/temperature cycle protocols, Packaging validation, Material Data Sheets
9 | Infection and Microbial contamination
9.1 |  | The medical devices and manufacturing processes shall be designed in such a way as to eliminate or to reduce as far as reasonably practicable and appropriate the risk of infection to any persons. The design shall: • allow easy handling, and, where necessary: • reduce as far as reasonably practicable and appropriate any microbial leakage from the medical device and/or microbial exposure during use, • if appropriate, minimises contamination of the medical device, or specimen where applicable, by the patient, user or other person, or contamination of the patient by the medical device, during its use. | Standards: ISO 14971 (Risk management) Others: Device design and composition, Risk management plan, Label information
9.2 |  | Where a medical device incorporates substances of biological origin, the risk of infection must be reduced as far as reasonably practicable and appropriate by selecting appropriate sources, donors and substances and by using, as appropriate, validated inactivation, conservation, test and control procedures. This may not apply to certain IVD medical device if the activity of the virus and other transmissible agent are integral to the intended purpose of the IVD medical device or when such elimination or inactivation process would compromise the performance of the IVD medical device. | Standards: ISO 22442 (Mds utilising animal tissues and their derivatives - Application of Risk Management), AAMI/ISO 14160 (Sterilisation of MDs incorporating materials of animal origin), ISO 10993 (Biocompatiility), ISO 14971 (Risk management) Others: Certificate of origin, Certificate of veterinary/abattoire inspection, Virsl validation studies, Biological safety studies
9.3 |  | Products incorporating non-viable tissues, cells and substances of animal origin falling within the definition of a medical device, shall originate from animals that have been subjected to veterinary controls and surveillance adapted to the intended purpose of the tissues. The product owner is required to retain information on the geographical origin of the animals. Processing, preservation, testing and handling of tissues, cells and substances of animal origin shall be carried out so as to provide optimal safety. In particular, safety with regard to viruses and other transmissible agents shall be addressed by implementation of validated methods of elimination or inactivation in the course of the manufacturing process. This may not apply to certain IVD medical device if the activity of the virus and other transmissible agent are integral to the intended purpose of the IVD medical device or when such elimination or inactivation process would compromise the performance of the IVD medical device.
9.4 |  | For products incorporating cells, tissues and derivatives of microbial or recombinant origin falling within the definition of a medical device, the selection of sources/donors, the processing, preservation, testing and handling of cells, tissues and derivatives of such origin shall be carried out so as to provide optimal safety. In particular, safety with regard to viruses and other transmissible agents shall be addressed by implementation of validated methods of elimination or inactivation in the course of the manufacturing process. This may not apply to certain IVD medical device if the activity of the virus and other transmissible agent are integral to the intended purpose of the IVD medical device or when such elimination or inactivation process would compromise the performance of the IVD medical device.
9.5 |  | For products incorporating non-viable human tissues, cells and substances falling within the definition of an IVD medical device, the selection of sources, donors and/or substances of human origin, the processing, preservation, testing and handling of tissues, cells and substances of such origin shall be carried out so as to provide optimal safety. In particular, safety with regard to viruses and other transmissible agents shall be addressed by implementation of validated methods of elimination or inactivation in the course of the manufacturing process. This may not apply to certain IVD medical device if the activity of the virus and other transmissible agent are integral to the intended purpose of the IVD medical device or when such elimination or inactivation process would compromise the performance of the IVD medical device.
9.6 |  | Medical devices labelled as having a special microbiological state shall be designed, manufactured and packed to ensure they remain so when placed on the market and remain so under the transport and storage conditions specified by the product owner. | Standards: EN 1174 (sterilisation - Estimation of population of micro organisms) Others: Bio-burden testing, packaging validation, stability testing (include transportation, storage and temperature cycle), validaiton of recommended storage conditions
9.7 |  | Medical devices delivered in a sterile state shall be designed, manufactured and packed to ensure that they remain sterile when placed on the market and remain sterile, under the transport and storage conditions indicated by the product owner. | Standards: Sterilisation standards (e.g. ISO 11137, AAMI TIR 27, ISO 11135, ISo 10993-7 (EO sterilisaiton Residuals), ISO 13408 (aseptic filtration), ISO 11607 (packaging of terminally sterilised MDs) Others: Sterilisation validation reports, Evidence of on-going validation, Bio-burden testing, SAL levels, EO residuals report (where applicable)
9.8 |  | Medical devices labelled either as sterile or as having a special microbiological state shall have been processed, manufactured and, if applicable, sterilised by appropriate, validated methods. | Standards: EN 556 (Requirements for MDs to be designated "STERILE") In addition, similar to EP 9.6. and/or 9.7 above
9.9 |  | Medical devices intended to be sterilised shall be manufactured in appropriately controlled (e.g. environmental) conditions. | Standards: EN 1174, USP and other relevant pharmacopeia standards Others: Bio-burden control and monitoring reports
9.1 |  | Packaging systems for non-sterile medical devices shall keep the product at the level of cleanliness stipulated and, if the medical devices are to be sterilised prior to use, minimise the risk of microbial contamination; the packaging system shall be suitable taking account of the method of sterilisation indicated by the product owner. The medical device shall be produced in appropriately controlled conditions. | Others: Packaging validation, Bio-burden control and monitoring, storage and handling recommendations
9.11 |  | The packaging and/or label of the medical device shall distinguish between identical or similar products placed on the market in both sterile and non-sterile condition. | Standards: ISO 13485 (QMS)
10 | Manufacturing and Environmental Properties
10.1 |  | If the medical device is intended for use in combination with other medical devices or equipment, the whole combination, including the connection system shall be safe and shall not impair the specified performance of the medical devices, or equipment with which it is used. Any restrictions on use applying to such combinations shall be indicated on the label and/or in the instructions for use. | Standards: ISO 60601 (Electrical safety collateral standards), ISO 14971 (Risk management), Compatibility standards (where available) Others: Device specifications, compatibility validation reports, Instructions for users
10.2 |  | Medical devices shall be designed and manufactured in such a way as to remove or reduce as far as reasonably practicable and appropriate: • the risk of injury, in connection with their physical features, including the volume/pressure ratio, dimensional and where appropriate ergonomic features; • risks connected with reasonably foreseeable external influences or environmental conditions, such as magnetic fields, external electrical and electromagnetic effects, electrostatic discharge, pressure, humidity, temperature or variations in pressure and acceleration; • the risks connected to their use in conjunction with materials, substances and gases with which they may come into contact during normal conditions of use; • the risks of accidental penetration of substances into the medical device; • the risk of incorrect identification of specimens; • the risks of reciprocal interference with other medical devices normally used in the investigations or for the treatment given; • risks arising where maintenance or calibration are not possible (as with implants), from ageing of materials used or loss of accuracy of any measuring or control mechanism. | Standards: ISO 60601 (Electrical safety and electromagnetic compatibility as applicable), ISO 14971 (Risk management), Product specific standards (where available) Others: Device specifications, Use environment and usability validation, Instructions for users
10.3 |  | Medical devices shall be designed and manufactured in such a way as to minimise the risks of fire or explosion during normal use and in single fault condition. Particular attention shall be paid to medical devices whose intended purpose includes exposure to or use in association with flammable substances or substances which could cause combustion. | Standards: ISO 60601 (Electrical safety as applicable), ISO 14971 (Risk management)
10.4 |  | Medical devices must be designed and manufactured in such a way as to facilitate the safe disposal of any waste substances. | Standards: ISO 14971 (Risk management) Others: Risk acceptace document, Material Safety Data Sheet
11 | Medical Device with a diagnostic or measuring function
11.1 |  | Medical devices with a measuring function shall be designed and manufactured in such a way as to provide sufficient accuracy, precision and stability for their intended purpose of the medical device. The limits of accuracy, precision and stability shall be indicated by the product owner. | Standards: Product specific standards (where available) Others: Performance verificaiton and validation, Information for use/device labelling
11.2 |  | Medical devices shall be designed and manufactured in such a way as to provide sufficient accuracy, precision and stability for their intended purpose, based on appropriate scientific and technical methods. In particular the design shall address the sensitivity, specificity, trueness, repeatability, reproducibility, control of known relevant interference and limits of detection, as appropriate. | Standards: Product specific standards (where available) Others: Performance verificaiton and validation, Information for use/device labelling
11.3 |  | Where the performance of medical devices depends on the use of calibrators and/or control materials, the traceability of values assigned to such calibrators and/or control materials shall be assured through a quality management system. | Standards: ISO 13485 (QMS), Product specific standards (where available) Others: Performance verificaiton and validation
11.4 |  | Any measurement, monitoring or display scale shall be designed and manufactured in line with ergonomic principles, taking into account of the intended purpose of the medical device. | Standards: ISO 62366 (Usability Engineering to MDs), Product specific standards (where available) Others: Usability studies
11.5 |  | Wherever possible values expressed numerically shall be in commonly accepted, standardised units, and understood by the users of the medical device. | Standards: ISO 15223 (labeling and information to be supplied with MDs), Product specific standards (where available) Others: Usability studies, Instructions for Use
12 | Protection against Radiation
12.1 |  | General
|  | Medical devices shall be designed and manufactured and packaged in such a way that exposure of patients, users and other persons to any emitted radiation shall be reduced as far as practicable and appropriate, compatible with the intended purpose, whilst not restricting the application of appropriate specified levels for therapeutic and diagnostic purposes | Standards: IEC 60601 (various parts as applicable), Product specific standards (where available) Others: Risk acceptace document
12.2 |  | Intended Radiation
12.2.1 |  | Where medical devices are designed to emit hazardous, or potentially hazardous, levels of visible and/or invisible radiation necessary for a specific medical purpose the benefit of which is considered to outweigh the risks inherent in the emission, it shall be possible for the user to control the emissions. Such medical devices shall be designed and manufactured to ensure reproducibility of relevant variable parameters within an acceptable tolerance. | Standards: IEC 60601 (e.g. For CT systems - IEC 60601-2-44; Radiation Protection in diagnostic X-ray - IEC 60601-1-3), Product specific standards (where available) Others: Test Reports, Risk acceptace document
12.2.2. |  | Where medical devices are intended to emit potentially hazardous, visible and/or invisible radiation, they shall be fitted, where practicable, with visual displays and/or audible warnings of such emissions. | Standards: IEC 60601 (e.g. For CT systems - IEC 60601-2-44), Product specific standards (where available) Others: Test Reports, Risk acceptace document
12.3 |  | Unintended Radiation
|  | Medical devices shall be designed and manufactured in such a way that exposure of patients, users and other persons to the emission of unintended, stray or scattered radiation is reduced as far as practicable and appropriate. | Standards: IEC 60601 (e.g. Radiation Protection in diagnostic X-ray - IEC 60601-1-3, General electrical safety), Product specific standards (where available) Others: Test Reports, Risk acceptace document
12.4 |  | Instructions for Use
|  | The operating instructions for medical devices emitting radiation shall give detailed information as to the nature of the emitted radiation, means of protecting the patient and the user and on ways of avoiding misuse and of eliminating the risks inherent in installation | Standards: IEC 60601 (e.g. Electrical safety and For CT systems - IEC 60601-2-44), ISO 15223 (labeling and information to be supplied with MDs), Product specific standards (where available) Others: Risk management report and Risk acceptace document
12.5 |  | Ionising Radiation
12.5.1 |  | Medical devices intended to emit ionising radiation shall be designed and manufactured in such a way as to ensure that, where practicable, the quantity, geometry and energy distribution (or quality) of radiation emitted can be varied and controlled taking into account the intended purpose. | Standards: IEC 60601 (e.g. For CT systems - IEC 60601-2-44; Radiation Protection in diagnostic X-ray - IEC 60601-1-3), IEC 61223 (e.g. Constancy tests - Imaging performance of CT - 61223-2-6; Acceptance tests - IEC 61223-3-5) Product specific standards (where available) Others: Test Reports, Risk acceptace document
12.5.2 |  | Medical devices emitting ionising radiation intended for diagnostic radiology shall be designed and manufactured in such a way as to achieve appropriate image and/or output quality for the intended medical purpose whilst minimising radiation exposure of the patient and user. | Standards: IEC 60601 (e.g. For CT systems - IEC 60601-2-44; Radiation Protection in diagnostic X-ray - IEC 60601-1-3), Product specific standards (where available) Others: Test Instructions and test reports
12.5.3 |  | Medical devices emitting ionising radiation, intended for therapeutic radiology shall be designed and manufactured in such a way as to enable reliable monitoring and control of the delivered dose, the beam type and energy and where appropriate the energy distribution of the radiation beam. | Standards: IEC 60601 (e.g. For gamma beam therapy equipment - IEC 60601-2-11), Product specific standards (where available) Others: Test Instructions and test reports
13 | Requirements for medical devices connected to or equipped with an energy source
13.1 |  | Medical devices incorporating electronic programmable systems, including software, shall be designed to ensure the repeatability, reliability and performance of these systems according to the intended purpose. In the event of a single fault condition in the system, appropriate means shall be adopted to eliminate or reduce as far as practicable and appropriate consequent risks. | Standards: IEC 60601 (all relevant parts of the standard), Product specific standards (where available), IEC 62304 (software), ISO 14971 (Risk management) Others: Device verification and validation, software validation report, Risk analysis,
13.2 |  | For medical devices which incorporate software or which are medical software in themselves, the software shall be validated according to the state of the art taking into account the principles of development lifecycle, risk management, validation and verification. | Standards: IEC 60601 (all relevant parts of the standard), Product specific standards (where available), IEC 62304 (software), ISO 14971 (Risk management) Others: Device verification and validation, software validation report, Risk analysis,
13.3 |  | Medical devices where the safety of the patients depends on an internal power supply shall be equipped with a means of determining the state of the power supply. | Standards: IEC 60601 (applicable parts), ISO 14971 (Risk management) Others: Device verification and validation
13.4 |  | Medical devices where the safety of the patients depends on an external power supply shall include an alarm system to signal any power failure. | Standards: IEC 60601 (IEC 60601-1-8 - alarms), ISO 14971 (Risk management) Others: Device verification and validation, Test Reports, Usability tests
13.5 |  | Medical devices intended to monitor one or more clinical parameters of a patient shall be equipped with appropriate alarm systems to alert the user of situations which could lead to death or severe deterioration of the patient's state of health.
13.6 |  | Medical devices shall be designed and manufactured in such a way as to reduce as far as practicable and appropriate the risks of creating electromagnetic interference which could impair the operation of this or other medical devices or equipment in the vicinity where the medical device is located. | Standards: IEC 60601 (EMC IEC 60601-1-2), ISO 14971 (Risk management) Others: Device verification and validation, Test Reports
13.7 |  | Medical devices shall be designed and manufactured in such a way as to provide an adequate level of intrinsic immunity to electromagnetic disturbance to enable them to operate as intended. | Standards: IEC 60601 (EMC IEC 60601-1-2), ISO 14971 (Risk management) Others: Device verification and validation, Test Reports
13.8 |  | Protection against electrical risks A medical device shall be designed and manufactured in a way that ensures that, as far as possible, a patient, or any other person is protected against the risk of accidental electric shock when it is installed and maintained as indicated by the product owner, is being used under normal conditions of use and in the event of a single fault condition. | Standards: IEC 60601 (applicable parts), ISO 14971 (Risk management) Others: Device verification and validation, Test reports
14 | Protection against mechanical risks
14.1 |  | Medical devices shall be designed and manufactured in such a way as to protect the patient and user against mechanical risks associated with the use of the medical device. | Standards: IEC 60601 (applicable parts), ISO 14971 (Risk management) Others: Device verification and validation, Test reports
14.2 |  | Medical devices shall be designed and manufactured in such a way as to reduce to the lowest practicable level the risks arising from vibration generated by the medical devices, taking account of technical progress and of the means available for limiting vibrations, particularly at source, unless the vibrations are part of the specified performance. | Standards: Product specific standards (where available), ISO 14971 (Risk management) Others: Device verification and validation, Vibration tests, Test reports
14.3 |  | Medical devices shall be designed and manufactured in such a way as to reduce to the lowest practicable level the risks arising from the noise emitted, taking account of technical progress and of the means available to reduce noise, particularly at source, unless the noise emitted is part of the specified performance. | Standards: Product specific standards (where available), IEC 60601 (applicable parts), ISO 14971 (Risk management) Others: Device verification and validation, Tests to evaluate noise level and other Test reports
14.4 |  | Terminals and connectors to the electricity, gas or hydraulic and pneumatic energy supplies which the user has to handle shall be designed and constructed in such a way as to minimise all possible risks.
14.5 |  | Accessible parts of the medical devices (excluding the parts or areas intended to supply heat or reach given temperatures) and their surroundings shall not attain potentially dangerous temperatures under normal use.
15 | Protection against the risks posed to the patient by supplied energy or substances
15.1 |  | Medical devices for supplying the patient with energy or substances shall be designed and constructed in such a way that the delivered rate and/or amount can be set and maintained accurately enough to guarantee the safety of the patient and of the user. | Standards: IEC 60601 (applicable parts), ISO 15223, EN 1041 (labeling and information to be supplied with MDs), ISO 14971 (Risk management) Others: Device verification and validation, Design and control mechanisms, Test reports, Risk analysis and acceptance document
15.2 |  | Medical devices shall be fitted with the means of preventing and/or indicating any inadequacies in the delivered rate and/or amount which could pose a danger. Medical devices shall incorporate suitable means to prevent, as far as possible, the accidental release of dangerous levels of energy from an energy and/or substance source.
15.3 |  | The function of the controls and indicators shall be clearly specified on the medical devices. Where a medical device bears instructions required for its operation or indicates operating or adjustment parameters by means of a visual system, such information shall be understandable to the user and, as appropriate, the patient.
16 | Active Implantable medical devices
16.1 |  | An active implantable medical device shall incorporate, display, emit or exhibit a code or unique characteristic that can be used to identify:- • the type of medical device; • the product owner of the medical device; and • the year of manufacture of the medical device. | Standards: ISO 15223, EN 1041 (labeling and information to be supplied with MDs), ISO 14971 (Risk management) Others: Labeling and Instructions for use, Implant card or other identification system for the user/patient
16.2 |  | The identifier shall be readable without the need for surgery to the person in whom the medical device is implanted.
17 | Protection against risks posed to the patient for medical devices for self-testing or self-administration
17.1 |  | Such medical devices shall be designed and manufactured in such a way that they perform appropriately for their intended purpose taking into account the skills and the means available to users and the influence resulting from variation that can reasonably be anticipated in user’s technique and environment. The information and instructions provided by the product owner shall be easy for the user to understand and apply. | Standards: ISO 62366 (Usability Engineering to MDs), Product specific standards (where available) Others: Usability studies
17.2 |  | Such medical devices shall be designed and manufactured in such a way as to reduce as far as practicable the risk of error in the handling of the medical device and, if applicable, the specimen, and also in the interpretation of results.
17.3 |  | Such medical devices shall, where reasonably possible, include a procedure by which the user can verify that, at the time of use, the medical device will perform as intended by the product owner.
18 | Information supplied by the product owner
|  | The following information shall be provided with a medical device, having regard to the training and knowledge of potential users of the medical device: • information identifying the medical device; • information identifying the product owner of the medical device; • information explaining how to use the medical device safely | Standards: IEC 60601 (applicable parts for active devices), ISO 15223, EN 1041 (labeling and information to be supplied with MDs), ISO 14971 (Risk management) Others: Labeling and Instrucitons for use, Usability tests
19 | Clinical Investigation
|  | Clinical investigations on human subjects shall be carried out in accordance with the spirit of the Helsinki Declaration. This includes every step in the clinical investigation from first consideration of the need and justification of the study to publication of the results. Clinical investigations on human subjects shall be carried out in accordance with the spirit of the Helsinki Declaration. This includes every step in the clinical investigation from first consideration of the need and justification of the study to publication of the results. | Standards: Product Specific Standards for performance evaluaiton (where available) Others: Clinical evidence and clinical evaluaiton report, Other verificaiton and validation reports

## Annex 2 Risk Class GMD
ANNEX 2 Risk Classification Rules for Medical Devices other than IVD Medical Devices
Section |  | Interpretative Notes - Examples (For discussion)
1. DEFINITIONS
ACTIVE MEDICAL DEVICE | Any medical device, operation of which depends on a source of electrical energy or any source of power other than that directly generated by the human body or gravity and which acts by converting this energy. Medical devices intended to transmit energy, substances or other elements between an active medical device and the patients, without any significant change, are not considered to be active medical devices. | Definition NOTE: Sources of power / energy - Includes chemical reaction (e.g. heat pack / cold pack by endo/exothermic reaction), potential / mechanical (e.g. spring loaded drug delivery system, sprays in pressurized canisters), radiation Converting this energy - Conversion may occur within the device (e.g. bone drills) , between device and tissue (e.g. electrocautery, ventilators), or within tissue (e.g. therapeutic HIFU)
| NOTE: Standalone software (to the extent it falls within the definition of a medical device) is deemed to be an active device. | Standalone - Software not embedded in another device; classified according to active rules
ACTIVE THERAPEUTIC DEVICE | Any active medical device, whether used alone or in combination with other medical devices, to support, modify, replace or restore biological functions or structures with a view to treatment or alleviation of an illness, injury or handicap. | Definition
ACTIVE DEVICE INTENDED FOR DIAGNOSIS | Any active medical device, whether used alone or in combination with other medical devices, to supply information for detecting, diagnosing, monitoring or to support in treating physiological conditions, states of health, illnesses or congenital deformities. | Definition
BODY ORIFICE | Any natural opening in the body, as well as the external surface of the eyeball, or any permanent artificial opening, such as a stoma or permanent tracheotomy. | Definition (i.e. ear, mouth, nose, eye, anus, urethra and vagina)
CENTRAL CIRCULATORY SYSTEM | For the purpose of this document, central circulatory system means the major internal blood vessels including the following: • arteriae pulmonales (pulmonary artery); • aorta ascendens (ascending aorta); • arteriae coronariae (coronary artery); • arteria carotis communis (common carotid artery); • arteria carotis externa (external carotid artery); • arteria carotis interna (internal carotid artery); • arteriae cerebrates (cerebella arteries); • truncus brachiocephalicus (brachiocephalic trunk); • venae cordis (cardiac veins); • venae pulmonales (pulmonary vein); • venae cava superior (superior vena cava); • venae cava inferior (inferior vena cava); • arcus aorta (aortic arch); • thoracica aorta (thoracic aorta); • abdominalis aorta (abdominal aorta); • arteriae ilica communis (common iliac arteries); • aorta descendens to the bifurcatio aortae. (descending aorta to the bifurcation of aorta) | Definition
CENTRAL NERVOUS SYSTEM | For the purpose of this document, central nervous system refers to the brain, meninges and spinal cord. | Definition
CONTINUOUS USE | in relation to a medical device, means • the uninterrupted use of the medical device, not including any temporary interruption of its use during a procedure or any temporary removal of the medical device for purposes such as cleaning or disinfection; or • the accumulated use of the medical device by replacing it immediately with another medical device of the same type, as intended by its product owner; | Definition (e.g. replaceable tracheostomy tubes)
DURATION OF USE - TRANSIENT | Normally intended for continuous use for less than 60 minutes | Definition
DURATION OF USE - SHORT TERM | Normally intended for continuous use for between 60 minutes and 30 days | Definition
DURATION OF USE - LONG TERM | Normally intended for continuous use for more than 30 days. | Definition
HARM | Physical injury or damage to the health of people or damage to property or the environment. | Definition
HAZARD | Potential source of harm. | Definition
IMMEDIATE DANGER | A situation where the patient is at risk of either losing life or an important physiological function if no immediate preventative measure is taken. | Definition
IMPLANTABLE MEDICAL DEVICE | Any medical device, including those that are partially or wholly absorbed, which is intended: - • to be totally introduced into the human body or, • to replace an epithelial surface or the surface of the eye, by surgical intervention which is intended to remain in place after the procedure. | Definition
| NOTE: Any medical device intended for partial introduction into the human body through surgical intervention and intended to remain in place after the procedure for at least 30 days is also considered an implantable medical device. | Definition
INVASIVE MEDICAL DEVICE | A medical device, which, in whole or in part, penetrates inside the body, either through a body orifice or through the surface of the body. | Definition
LIFE SUPPORTING OR LIFE SUSTAINING | A medical device that is essential to, or that yields information that is essential to, the restoration or continuation of a bodily function important to the continuation of human life. | Definition
REUSABLE SURGICAL INSTRUMENT | Instrument intended for surgical use by cutting, drilling, sawing, scratching, scraping, clamping, retracting, clipping or other surgical procedures, without connection to any active medical device and which are intended by the product owner to be reused after appropriate procedures for cleaning and/or sterilisation have been carried out. | Definition
RISK | Combination of the probability of occurrence of harm and the severity of that harm. | Definition
SURGICALLY INVASIVE MEDICAL DEVICE | An invasive medical device that penetrates inside the body through the surface of the body, with the aid or in the context of a surgical operation. | Definition
2. RISK CLASSIFICATION FOR MEDICAL DEVICES OTHER THAN IVD MEDICAL DEVICES
Non-invasive
Rule 1 | All non-invasive medical devices which come into contact with injured skin:
| are in Class A if they are intended to be used as a mechanical barrier, for compression or for absorption of exudates only, i.e. they heal by primary intent; | Injured skin where dermis is not breached e.g. plaster
| are in Class B if they are intended to be used principally with wounds which have breached the dermis, including medical devices principally intended to manage the microenvironment of a wound. | manage microenvironent - e.g. control level of moisture / oxygen or other parameters around the wound to facilitate healing process
| Unless they are intended to be used principally with wounds which have breached the dermis and can only heal by secondary intent, in which case they are in Class C. | Full / partial thickness wounds, e.g. dressings for chronic ulcerated wounds, dressings for severe burns.
Rule 2 | All non-invasive medical devices intended for channelling or storing • body liquids or tissues, • liquids or • gases for the purpose of eventual infusion, administration or introduction into the body are in Class A, | channelling / storing of non-body liquids or gases. E.g. administration sets for gravity infusion; syringes without needles
| Unless they may be connected to an active medical device in Class B or a higher class, in which case they are Class B; | e.g. administration sets for infusion pumps
| Unless they are intended for use of • channeling blood, or • storing or channeling other body liquids, or • for storing organs, parts of organs or body tissues, in which case they are Class B. | e.g. tubings for transfusion, organ storage containers
Unless they are blood bags, in which case they are Class C.
Rule 3 | All non-invasive medical devices intended for modifying the biological or chemical composition of • blood, • other body liquids, or • other liquids intended for infusion into the body are in Class C, | e.g. hemodialysis concentrates
| Unless the treatment consists of filtration, centrifuging or exchanges of gas or of heat, in which case they are in Class B. | Modifying composition by "mechanical" methods; by adding / removal of gas; modifying by changing temperature. E.g. particulate filters in extracorporeal circulation system.
Rule 4 | All other non-invasive medical devices are in Class A. | Fallback rule, all other non-invasive devices not covered by rules 1-3 e.g. urine collection bottles; non-invasive electrodes
Invasive
Rule 5 | All invasive medical devices with respect to body orifices (other than those which are surgically invasive) and which: • are not intended for connection to an active medical device, or • are intended for connection to a Class A medical device only. | "body orifice" - i.e. invasive through ear, mouth, nose, eye, anus, urethra or vagina
| - are in Class A if they are intended for transient use; | "transient": <60 mins e.g. examination gloves; handheld dental mirror; vaginal speculum; manual nasal aspirator; eye irrigation kit (w/o solution)
| Unless they are intended by its product owner for use on the external surface of any eyeball; or it is liable to be absorbed by the mucous membrane, in which case they are in Class B. | mucous membranes through body orifices - mouth, respiratory system, eyelids e.g. eye drops, eye irrigation solution; artificial tears; artificial saliva / saliva substitute
| - are in Class B if they are intended for short-term use; | "short-term": 60 mins - 30 days e.g. urinary catheters for short-term use; tracheal tubes; contact lens
| Unless they are intended for short-term use in the oral cavity as far as the pharynx, in an ear canal up to the ear drum or in a nasal cavity, in which case they are in Class A, | e.g. intranasal splint, dressing for nose bleeds
| - are in Class C if they are intended for long-term use; | "long-term": > 30 days e.g. urethral stent
| Unless they are intended for long-term use in the oral cavity as far as the pharynx, in an ear canal up to the ear-drum or in a nasal cavity and are not liable to be absorbed by the mucous membrane, in which case they are in Class B. | e.g. fixed dental prosthesis
| All invasive medical devices with respect to body orifices (other than those which are surgically invasive) that are intended to be connected to an active medical device in Class B or a higher class, are in Class B. | Regardless of duration of use e.g. tracheal tubes connected to a ventilator; suction catheters for stomach drainage; dental aspirator tips.
Rule 6 | All surgically invasive medical devices intended for transient use are in Class B, | "transient": <60 mins
| Unless they are reusable surgical instruments, in which case they are in Class A; or | e.g. manual drill bit
| Unless intended to supply energy in the form of ionising radiation, in which case they are in Class C; or | e.g. catheters containing or incorporating sealed radioisotopes
| Unless intended to have a biological effect or be wholly or mainly absorbed, in which case they are in Class C; or | "biological effect" - intentionally induces/changes a reaction from the body/tissue; does not include unintentional / adverse effects e.g. insufflation gases
| Unless intended to administer medicinal products by means of a delivery system, if this is done in a manner that is potentially hazardous taking account of the mode of application, in which they are in Class C; or | Administration = storage and/or influencing the rate/volume of medicine delivered not just channelling. e.g. insulin pen
| Unless they are intended specifically for use in direct contact with the central nervous system, in which case they are in Class D; or | e.g. brain spatulas
| Unless intended specifically to diagnose, monitor or correct a defect of the heart or of the central circulatory system through direct contact with these parts of the body, in which case they are in Class D. | e.g. angioplasty balloon catheters, stent delivery catheters/systems, including related guidewires and related introducers
Rule 7 | All surgically invasive medical devices intended for short-term use are in Class B, | "short-term": 60 mins - 30 days e.g. infusion cannulae, non-absorbable skin closure devices
| Unless they are intended to administer medicinal products, in which case they are in Class C; or | Administration = storage and/or influencing the rate/volume of medicine delivered not just channelling.
Unless they are intended to undergo chemical change in the body (except if the medical devices are placed in the teeth), in which case they are in Class C; or
| Unless they are intended to supply energy in the form or ionising radiation, in which case they are in Class C; or | e.g. brachytherapy devices
| Unless they are intended to have a biological effect or to be wholly or mainly absorbed, in which case they are in Class D; or | "biological effect" - intentionally induces/changes a reaction from the body/tissue; does not include unintentional / adverse effects e.g. short-term absorbable sutures, absorbable meshes
| Unless they are intended specifically for use in direct contact with the central nervous system, in which case they are in Class D; | e.g. neurological catheter.
| Unless they are intended specifically to diagnose, monitor or correct a defect of the heart or of the central circulatory system through direct contact with these parts of the body, in which case they are in Class D. | e.g. temporary pacemaker leads
Rule 8 | All implantable medical devices, and long-term surgically invasive medical devices, are in Class C, | "long-term": >30 days e.g. orthopedic implants, non-absorbable sutures
| Unless they are intended to be placed into the teeth, in which case they are in Class B; or | e.g. bridges, crowns, dental filling
| Unless they are intended to be used in direct contact with the heart, the central circulatory system or the central nervous system, in which case they are in Class D; or | e.g. prosthetic heart valves, coronary stent
Unless they are intended to be life supporting or life sustaining, in which case they are in Class D; or
| Unless they are intended to be active implantable medical devices, in which case they are Class D; or | e.g. sub-cutaneous ICD, implantable cardiac monitor
| Unless they are intended to have a biological effect or to be wholly or mainly absorbed, in which case they are in Class D; or | "biological effect" - intentionally induces/changes a reaction from the body/tissue; does not include unintentional / adverse effects e.g. long-term absorbable sutures. implants with bioactive surface coatings
| Unless they are intended to administer medicinal products, in which case they are in Class D; or | e.g. implantable infusion pump
Unless they are intended to undergo chemical change in the body (except if the medical devices are placed in the teeth), in which case they are in Class D; or
| Unless they are breast implants, in which case they are in Class D. | Include tissue expanders
Active
Rule 9(i) | All active therapeutic medical devices intended to administer or exchange energy are in Class B, | e.g. TENS (Transcutaneous Electro-Neuro Stimulator), hearing aids, rehabilitation machines
| Unless their characteristics are such that they may administer or exchange energy to or from the human body in a potentially hazardous way, including ionising radiation, taking account of the nature, the density and site of application of the energy, in which case they are in Class C. | All devices administering ionising radiation for therapeutic purpose are classifed under this rule. NOTE: ‘potentially hazardous’ refers to the type of technology involved and the intended application. E.g. electrosurgical generators, HIFU e.g. lung ventilators, baby incubators, external defibrillators
Rule 9(ii) | All active medical devices intended to control or monitor the performance of active therapeutic medical devices in Class C, or intended directly to influence the performance of such medical devices, are in Class C. | e.g. sound processor controlling Class D cochlear implant
Rule 10(i) | Active medical devices intended for diagnosis are in Class B:
| - if they are intended to supply energy which will be absorbed by the human body (except for medical devices used solely to illuminate the patient's body, with light in the visible or near infra-red spectrum, in which case they are Class A), or | Diagnostic MDs which supply all types of energy to be absorbed by the body are covered by this rule, with the exception of light intended for illumination (exception under this rule) and ionising radiation (covered under rule 10(ii)).
| - if they are intended to image in vivo distribution of radiopharmaceuticals, or | e.g. gamma cameras
| - if they are intended to allow direct diagnosis or monitoring of vital physiological processes, | "vital physiological processes" - e.g. body temperature, heart rate, ECG, respiration, blood pressure For non-critical situations, e.g. routine check-up or self-monitoring e.g. thermometers, blood pressure monitors
| Unless they are specifically intended for: • monitoring of vital physiological parameters, where the nature of variations is such that it could result in immediate danger to the patient, for instance variations in cardiac performance, respiration, activity of central nervous system, or • diagnosing in clinical situations where the patient is in immediate danger, in which case they are in Class C. | "specifically intended" - Intended Use / Labelling claims made to indicate the criticality of parameter measured or device use setting, e.g.devices used for patient monitoring in ICU, intended use typically states that it is to be used in emergency care, critical/ acute care settings and such devices typically contain alarms to alert clinicians when vital physiological parameters go beyond the accepted clinical range e.g. patient monitors for ICU For critical situations, e.g. intensive care setting
Rule 10(ii) | Active medical devices intended to emit ionising radiation and intended for diagnostic and/or interventional radiology, including medical devices which control or monitor such medical devices, or those which directly influence their performance, are in Class C. | e.g. Diagnostic x-rays and devices (including software) that control, monitor or incluence the emission of x-ray
Rule 11 | All active medical devices intended to administer and/or remove medicinal products, body liquids or other substances to or from the body are in Class B, | e.g. suction equipment, feeding pumps
| Unless this is done in a manner that is potentially hazardous, taking account of the nature of the substances involved, of the part of the body concerned and of the mode and route of administration or removal, in which case they are in Class C. | e.g. infusion pumps, anaesthesia equipment, dialysis equipment
Rule 12 | All other active medical devices are in Class A. | Fallback rule, all other active devices not covered by rules 9-11. e.g. examination lamps, surgical microscopes, powered hospital beds, wheelchairs, dental curing lights.
Additional Rules
Rule 13 | All medical devices incorporating, as an integral part, a substance which, if used separately, can be considered to be a medicinal product (as defined by the Member State), and which is liable to act on the human body with action ancillary to that of the medical devices, are in Class D. | Rule for Device-Drug with a secondary pharmaceutical mode of action and intended use. "incorporating, as an integral part" - device and medicinal product supplied together and combined before or at the time of use/administration e.g. drug-eluting stent, dermal fillers with lidocaine, silver dressings (with antimicrobial claims made for silver)
Rule 14 | All medical devices manufactured from or incorporating • animal cells, tissues and/or derivatives thereof, rendered non-viable, or • cells, tissues and/or derivatives of microbial or recombinant origin are Class D, | e.g. porcine heart valves; catgut sutures; dermal fillers based on hyaluronic acid derived from bacterial fermentation processes
| Unless such medical devices are manufactured from or incorporate non-viable animal tissues or their derivatives that come in contact with intact skin only, where they are in Class A. | e.g. leather components of orthopaedic appliances
Rule 15 | All medical devices intended specifically to be used for sterilising medical devices, or disinfecting as the end point of processing, are in Class C. | "end point of processing" - final processing step before use by, or on, a patient. "sterilising" / "disinfecting" - device serves to kill or remove microorganisms NOTE: Rule does not apply for basic cleaning
| Unless they are intended for disinfecting medical devices prior to end point sterilisation or higher level disinfection, in which case they are in Class B; or | Claim sensitive
| Unless they are intended specifically to be used for disinfecting, cleaning, rinsing or, when appropriate, hydrating contact lenses, in which case they are in Class C. | e.g. contact lens multi-purpose solutions, eyedrops for hydrating contact lens
Rule 16 | All medical devices used for contraception or the prevention of the transmission of sexually transmitted diseases are in Class C, | e.g. condoms, contraceptive diaphragms
| Unless they are implantable or long-term invasive medical devices, in which case they are in Class D. | e.g. IUD (intrauterine device)

## Annex 3 Risk Class IVD
ANNEX 3: Risk Classification Rules for IVD Medical Devices
Section |  | Interpretative Notes - Examples (For discussion)
1. DEFINITIONS
EXAMINATION | Set of operations having the object of determining the value of a property. NOTE: Examination of an analyte in a biological sample is commonly referred to as a test, assay or analysis. | Definition
INSTRUMENT | Equipment or apparatus intended by the product owner to be used as IVD medical device. | Definition. NOTE: instrument software included under this definition.
IVD MEDICAL DEVICE FOR SELF-TESTING | Any IVD medical device intended by the product owner for use by lay persons. | Definition
LAY PERSON | Any individual who does not have formal training in a relevant field or discipline. | Definition
NEAR PATIENT TESTING | Any testing performed outside a laboratory environment by a healthcare professional not necessarily a laboratory professional, generally near to, or at the side of, the patient. Also known as Point-of-Care (POC). | Definition. Also known as bedside testing.
REAGENT | Any chemical, biological or immunological components, solutions or preparations intended by the product owner to be used as IVD medical devices. | Definition
SELF-TESTING | Testing performed by lay persons. | Definition
SPECIMEN RECEPTACLE | An IVD medical device, whether vacuum-type or not, specifically intended by their product owner for the primary containment of specimens derived from the human body. | Definition
TRANSMISSIBLE AGENT | An agent capable of being transmitted to a person, as a communicable, infectious or contagious disease. | Definition
TRANSMISSION | The conveyance of disease to a person. | Definition
2. CLASSIFICATION RULES FOR IVD MEDICAL DEVICES
Rule 1 | IVD medical devices intended for the following purposes are classified as Class D: • medical devices intended to be used to detect the presence of, or exposure to, a transmissible agent in blood, blood components, blood derivatives, cells, tissues or organs in order to assess their suitability for transfusion or transplantation, or • medical devices intended to be used to detect the presence of, or exposure to, a transmissible agent that causes a life-threatening, often incurable, disease with a high risk of propagation. | Scope of rule 1
Rationale for rule 1 | The application of this rule as defined above should be in accordance with the rationale that follows: IVD medical devices in this Class are intended to be used to ensure the safety of blood and blood components for transfusion and/or cells, tissues and organs for transplantation. In most cases, the result of the test is the major determinant as to whether the donation/product will be used. Serious diseases are those that result in death or long-term disability, which are often incurable or require major therapeutic interventions and where an accurate diagnosis is vital to mitigate the public health impact of the condition. | Class D devices represent high individual risk and high public health risk
Examples for rule 1 | Tests to detect infection by HIV, HCV, HBV, HTLV. This Rule applies to first-line assays, confirmatory assays and supplemental assays. | More specific examples: HIV blood donor screening, HIV blood diagnostic, hepatitis B and C antigen kits for diagnosis, Kits for identification of virulent strains like H1N1.
Rule 2 | IVD medical devices intended to be used for blood grouping, or tissue typing to ensure the immunological compatibility of blood, blood components, cells, tissue or organs that are intended for transfusion or transplantation, are classified as Class C, except for ABO system [A (ABO1), B (ABO2), AB (ABO3)], rhesus system [RH1 (D), RH2 (C), RH3 (E), RH4 (c), RH5 (e)], Kell system [Kel1 (K)], Kidd system [JK1 (Jka), JK2 (Jkb)] and Duffy system [FY1 (Fya), FY2 (Fyb)] determination which are classified as Class D. | Scope of rule 2; with distinction between Class C and Class D
Rationale for Rule 2 | The application of this rule as defined above should be in accordance with the following rationale: A high individual risk, where an erroneous result would put the patient in an imminent life-threatening situation places the medical device into Class D. The rule divides blood-grouping IVD medical devices into two subsets, Class C or D, depending on the nature of the blood group antigen the IVD medical device is designed to detect, and its importance in a transfusion setting. | Class C devices present high inidvidual risk and/ or moderate public health risk whereas Class D devices represent high individual risk and high public health risk
Examples for Rule 2 | HLA, Duffy system (other Duffy systems except those listed in the rule as Class D) are in Class C.
Rule 3 | IVD medical devices are classified as Class C if they are intended for use: • in detecting the presence of, or exposure to, a sexually transmitted agent (e.g. Sexually transmitted diseases, such as Chlamydia trachomatis, Neisseria gonorrhoeae). | Scope of rule 3
| • in detecting the presence in cerebrospinal fluid or blood of an infectious agent with a risk of limited propagation (e.g. Neisseria meningitidis or Cryptococcus neoformans). | Additional examples: rubella, pertussis identification kits
| • in detecting the presence of an infectious agent where there is a significant risk that an erroneous result would cause death or severe disability to the individual or fetus being tested (e.g. diagnostic assay for CMV, Chlamydia pneumoniae, Methycillin Resistant Staphylococcus aureus).• in pre-natal screening of women in order to determine their immune status towards transmissible agents (e.g. Immune status tests for Rubella or Toxoplasmosis). | Additional Examples: VRE (vancomycin-resistant enterococcus)
• in determining infective disease status or immune status, and where there is a risk that an erroneous result will lead to a patient management decision resulting in an imminent life-threatening situation for the patient (e.g. Enteroviruses, CMV and HSV in transplant patients).
| • in screening for selection of patients for selective therapy and management, or for disease staging, or in the diagnosis of cancer (e.g. personalised medicine). NOTE: Those IVD medical devices where the therapy decision would usually be made only after further investigation and those used for monitoring would fall into class B under rule 6. | Markers must be the sole/ key determinant in these conditions
• in human genetic testing (e.g. Huntington’s Disease, Cystic Fibrosis).
| • to monitor levels of medicines, substances or biological components, when there is a risk that an erroneous result will lead to a patient management decision resulting in an immediate life-threatening situation for the patient (e.g. Cardiac markers, cyclosporin, prothrombin time testing). | Additional examples of Class C kits: kits that measure the level of narrow therapeutic index drugs in serum levels (aminoglycosides, vancomycin, theophylline, barbiturates, methotrexate). Those for drugs with a wide therapeutic index are Class B (i.e. benzodiazepines) Examples of Class C cardiac markers kits: CK-MB in diagnosis of myocardial infarction
| • in the management of patients suffering from a life-threatening infectious disease (e.g. HCV viral load, HIV Viral Load and HIV and HCV geno- and subtyping). | Example of intended use for a Class C HCV kit: intendedfor use as an aid in the management of HCV-infected patients undergoing antiviral therapy. The assay is not for screening blood, plasma, serum or tissue donors for HCV,or to be used as a diagnostic test to confirm the presence of HCV infection.
| • in screening for congenital disorders in the fetus (e.g. Spina Bifida or Down Syndrome). | Trisomy 21 detection kits
Rationale for Rule 3 | The application of this rule as defined above should be in accordance with the rationale for this rule which is as follows: IVD medical devices in this Class present a moderate public health risk, or a high individual risk, where an erroneous result would put the patient in an imminent life-threatening situation, or would have a major negative impact on outcome. The IVD medical devices provide the critical, or sole, determinant for the correct diagnosis. They may also present a high individual risk because of the stress and anxiety resulting from the information and the nature of the possible follow-up measures. | Summary: moderate public health OR high individual public risk
Rule 4 | IVD medical devices intended for self-testing are classified as Class C, except those medical devices from which the result is not determining a medically critical status, or is preliminary and requires follow-up with the appropriate laboratory test in which case they are Class B.IVD medical devices intended for blood gases and blood glucose determinations for near-patient testing would be Class C. Other IVD medical devices that are intended for near patient should be classified in their own right using the classification rules. | Scope of rule 4
Rationale for Rule 4 | The application of this rule as defined above should be in accordance with the rationale for this rule which is as follows: In general, these IVD medical devices are used by individuals with no technical expertise and thus the labelling and instructions for use are critical to the proper outcome of the test.
Example for Self-testing Class C | Blood glucose monitoring, | Additional example: PT/ INR kit used to monitor the clotting time of blood (for use to monitor patients who are taking anticoagulant medications like warfarin in outpatient settings
Example for Self-testing Class B | Pregnancy self test, Fertility testing, Urine test strip. | Additional example: Blood gas testing
Rule 5 | The following IVD medical devices are classified as Class A: • reagents or other articles that possess specific characteristics, intended by the product owner to make them suitable for in-vitro diagnostic procedures related to a specific examination. • instruments intended by the product owner specifically to be used for in-vitro diagnostic procedures. • specimen receptacles. NOTE: Any product for general laboratory use not manufactured, sold or represented for use in specified in-vitro diagnostic applications are deemed to not be IVD medical devices. | Scope of rule 5. Instrument software is covered under this rule.
Rationale for Rule 5 | The application of this rule as defined above should be in accordance with the rationale for this rule which is as follows: These IVD medical devices present a low individual risk and no or minimal public health risk. | Low individual and public health risk
Examples for Rule 5 | Selective/differential microbiological media (excluding the dehydrated powders which are considered not to be a finished IVD medical device), identification kits for cultured microorganisms, wash solutions, instruments and plain urine cup. NOTE: The performance of software or an instrument that is specifically required to perform a particular test will be assessed at the same time as the test kit. NOTE: The interdependence of the instrument and the test methodology prevents the instrument from being assessed separately, even though the instrument itself is still classified as Class A. | Other examples: Diluents for reagents, consumables for instruments (cleaning solutions), reagent curvettes, software used for recording/ trending of results but do not provide diagnostic evaluation of results Additional Class A microbiology examples: Culture media that only identify a broad family/ group of microorganisms (e.g. Enterobacteria, anaerobic bacteria, legionella). Contrasted with Class C microbiological examples: Selective culture media that identifies legionella pneumonia/ salmonella typhi strain that cause typhoid fever.
Rule 6 | IVD medical devices not covered in Rules 1 through 5 are classified as Class B. | Scope of rule 6
Rationale for Rule 6 | The application of this rule as defined above should be in accordance with the rationale for this rule which is as follows: These IVD medical devices present a moderate individual risk as they are not likely to lead to an erroneous result that would cause death or severe disability, have a major negative impact on patient outcome or put the individual in immediate danger. The IVD medical devices give results that are usually one of several determinants. If the test result is the sole determinant however other information is available, such as presenting signs and symptoms or other clinical information that may guide a physician, such that classification into Class B may be justified. Other appropriate controls may also be in place to validate the results. This Class also includes those IVD medical devices that present a low public health risk because they detect infectious agents that are not easily propagated in a population. | Summarized: Devices fall under this rule when they present (i) moderate individual risk and/or (ii) low public health risk. For (i) First Example: These devices are not the marker for acute/ critical conditions based on their IFU claims and clinical guidelines. For example: using ESR for sepsis/ infection as it is a non-specific marker of inflammation. Examples of Class C kits for sepsis, procalcitonin and CRP Second Example: These devices are not a key or independent marker for a disease/ condition based on their IFU claims and clinical guidelines. Examples of Class B: kits of staining monoclonal antibodies that do not claim to a specific cancer type, OR antigens that are identified are not tied to a particular cancer type (e.g. p53). For Class C kit staining antibody kit, the identified antigen should be the hallmark of a specific cancer (e.g. her-2 and breast cancer, CA-125 and ovarian cancer, PSA and prostate cancer). (ii) Low public risk: Examples of kits that detect infectious agents that are not easily propagated or agents that pose a lower level of public health concern are, influenza A and B kits that are not able to identify specific strains of public concern like H1N1.
Examples for Rule 6 | Blood gases, H. pylori and physiological markers such as hormones, vitamins, enzymes, metabolic markers, specific IgE assays and celiac disease markers.
Rule 7 | IVD medical devices that are controls without a quantitative or qualitative assigned value will be classified as Class B. | Scope of rule 7
Rationale for Rule 7 | For such controls, the user, not the product owner, assigns the qualitative or quantitative value. | Such devices are hardly seen. Most controls are either qualitative (i.e. positive or negative controls) or semi-qualitative (i.e. low positive, medium positive and high positive) or quantitative

## Annex 4 CSDT
Refer to CSDT guidelines under MDPWG TC

## Annex 5 PMAS
Refer to Guidelines on Post Market Surveillance under AMDTC

## Annex 6 DoC
ARTICLE 8 TECHNICAL DOCUMENTS FOR MEDICAL DEVICES
Member States shall undertake appropriate measures to adopt and implement Annex 6 - Harmonized set of elements for a Product Owner’s or Physical Manufacturer’s Declaration of Conformity (DoC).
ANNEX 6 Components Elements of a Product Owner's or Physical Manufacturer's Declaration of Conformity (DOC)
Interpretation/ Comments (For discussion)
1 | COMPONENTS OF A DECLARATION OF CONFORMITY
The DOC shall contain the following information:
• an attestation that each medical device that is subject to the declaration • complies with the applicable Essential Principles for Safety and Performance, and • has been classified according to the classification rules;
• information sufficient to identify the medical device/s to which the DOC applies;
• the risk class allocated to the medical device/s after following the guidance found in Principles of Medical Device Classification;
• the date from which the DOC is valid;
• the name and address of the product owner and physical manufacturer;
• quality management standards;
• medical device standards (product standards1);
• the name, position and signature of the responsible person who has been authorised to complete the DOC upon the product owner’s behalf.
2 | RESPONSIBILITY FOR PREPARING THE DECLARATION OF CONFORMITY
The product owner or physical manufacturer of the medical device is responsible for preparing and signing the DOC. A copy of the signed and dated DOC should be submitted as part of product registration. The original signed copy of the DOC should be made available to the Regulatory Authority of a Member State upon request. A Member State may impose additional measures (e.g. legalisation or notarisation) to be undertaken to ensure the authenticity of a DOC submitted to the Regulatory Authority of that Member State.