PRELIMINARY LIST OF BORDERLINE/ MD OR NON MD CLASSIFICATION OF MEDICAL DEVICE 
Harmonized
	NO.
	PRODUCT NAME
	DESCRIPTION
	Borderline with/ MD or non MD
	AMDTC Decision

	5th AMDTC

	1.
	Micropipette 
	The standard laboratory equipment used to measure and transfer small volumes of liquids
	Medical Device or non Medical Device 
	Non Medical Device , unless Product Owner state it for IVD purposes



	2.
	Culture Media or Petri dish with culture media 
	Any material prepared for the growth of bacteria or other microorganism in a laboratory.
	Medical Device  or non Medical Device 
	Non Medical Device , unless Product Owner state it for IVD and IVF purposes



	3.
	Disinfectant/ antiseptic for dental
	Disinfectants or antiseptics for dental that contain materials: chlorhexidine gluconate, iodine, formaldehyde (?), Sodium hypochloride


	Medical Device  Borderline with Medicine
	Depends on the mode of action of the active ingredient. If it is physical mode of action then it is a Medical Device . But if it is pharmacological action, the it is classified as medicine.

15th AMDTC Meeting agreed to remove this item.

	4.
	EDTA Solution
	EDTA Solution buffered to neutral pH used for preparation of root canals by rapid cleansing and enlarging of root canals


	Medical Device  Borderline with Medicine
	Depends on the mode of action of the active ingredient. If it is physical mode of action then it is a Medical Device . But if it is pharmacological action, the it is classified as medicine.



	5.
	Mentholated bandage
	Mentholated bandage is a bandage containing menthol usually used to relieve pain
	Medical Device  Borderline with Medicine
	Non Medical Device 


	6.
	Eye drops
	Eye drops are saline – containing drops used as an ocular route to administer. Depending on the condition being treated, they may contain steroids, antihistamines, sympathomimetics, beta receptor blockers, parasympathomimetics, parasympatholytics, prostaglandins, non stroisal anti-inflammatory drugs (NSAIDs), antibiotics, antifungal or topical anesthetics. Eye drops sometimes do not have medications in them and are only lubricating and tear-replacing solution.
	Medical Device  Borderline with Medicine
	Medical Device  if specifically intended to be used for lubricating, disinfecting, cleaning, rinsing or, when 

appropriate, hydrating contact lenses and tear-replacing solution



	7.
	Artificial tears
	Artificial tears are lubricant eye drops used to treat the dryness and irritation associated with deficient tear production in keratoconjunctivitis sicca (dry eyes). They are also used to moisten contact lenses and in eye examinations.


	Medical Device  Borderline with Medicine
	Medical Device  if there are no active ingredients and specifically intended to be used for lubricating and tear-replacing solution

	8.
	Lotion (example is for prevention of bed sores)
	Moisturizing cream or skin protectant to treat skin gently to help prevent bed sores
	Medical Device  Borderline with Medicine
	Depends on the mode of action of the active ingredient. If it is physical mode of action then it is a Medical Device . But if it is pharmacological action, the it is classified as medicine.

	9.
	Personal Combinational Stimulator
	A therapeutic tool which can generate vibration waves, can help reduce and relieve aches and fatigue on the body.
	Medical Device  or non Medical Device 
	Medical Device 

	10.
	Bed Matres
	heat therapy device with infrared remote, can help reduce and relieve aches and fatigue on the body.
	Medical Device  or non Medical Device 
	Medical Device 

	11.
	Needle for tattooing
	Used in the context of cultural or individual
	Borderline with/ MD or non MD
	Non Medical Device 

	12.
	Needle used in the acupuncture
	Thin needles are inserted through the skin in which practitioners stimulate specific points on the body.
	Borderline with/ MD or non MD
	Medical Device 


	13.
	Laser
	Used to treat skin disorders and providing tangible cosmetic alterations to the skin
	Borderline with/ MD or non MD
	Medical Device  if there are any medical claim.

Non Medical Device , if for aesthetic/cosmetic purposes.

	14.
	Liposuction machine
	Intended to remove fats from the body using suction
	Borderline with/ MD or non MD
	Medical Device  



	15.
	Admission Kit
	Intended use during admission in the hospital. Contains, towel, toothbrush, glass, measuring cup, medicine cup, etc.
	Borderline with/ MD or consumer products
	Non Medical Device 

	16.
	Gadgets (ex. Cellphone) with health and medical applications
	Cellphones are being used in the telecommunication with applications in health monitoring such as heart rate
	Borderline with/ MD or non MD
	Non Medical Device 
The Application/Software is Medical Device 

	17.
	Tissue replacement of biological origin
	Regeneration and remodeling of tissue in vivo for the purpose of repairing, replacing, maintaining, or enhancing organ function, and the engineering and growing of functional tissue substitues in vitro for implantation in vivo as a biological substitue for damaged or diseased tissues and organs. It is either (1) manufactured from human tissue;  (2) manufactured from animal tissue, or  (3) blood substitues and expander.
	Borderline with/ MD or drugs
	Non Medical Device 
Medical Device , if there are no biological/pharmacological/metabolical action, non viable and manufactured from animal tissue.

	18.
	Laboratory equipments (microscope, centrifuge, etc.)
	Intended use in the preparation and analysis of specimen 
	Borderline with MD or non Md
	Non Medical Device , if General Laboratory Purposes

Medical Device , if use for IVD, IVF

	19.
	Laboratory products (pipettes, graduated cylinder, slides, beakers, etc.)
	Intended use for the preparation of specimen, aspirating and dispensing specific volumes in the microlitre scale
	Borderline with MD or not MD
	Non Medical Device , if General Laboratory Purposes

Medical Device , if use for IVD, IVF

	20.
	Eye drops
	Intended use in the alleviation of soreness, discomfort or irritation caused by environmental factors (such as dust, heat, smoke, etc.)
	Borderline with/ MD or drugs
	Medical Device  if specifically intended to be used for lubricating, disinfecting, cleaning, rinsing or, when 

appropriate, hydrating contact lenses and tear-replacing solution

	21.
	Washing solutions used for pathogenic microorganisms
	Intended use for disinfection.
	Borderline with/ MD or non MD
	Medical Device , if intended for disinfecting Medical Device s

	22.
	Hand Gloves
	The intended use is to reduce bacterial transfer
	Borderline with/ MD or non MD
	Medical Device  if with medical purpose

	23.
	Mobile application for the communication between patient and caregivers while giving birth
	This application allows storing data on each moment of contraction during delivery.
	Borderline with/ MD or non MD
	Non Medical Device 

	24.
	Software for information management and patient monitoring
	Intended to route and store Medical Device data
	Borderline with/ MD or non MD
	Medical Device  if used to monitor the patient data.

Non MD if used only for data storage.

	25.
	Wound treatment products
	Intended use to treat wounds.
	Borderline with/ MD or drugs
	Medical Device 

	26.
	Lubricant
	Intended use to reduce friction or to ease penetration
	Borderline with/ MD or non MD
	Medical Device 

	27.
	Medical Oxygen gas
	Indication as gases used for medical purposes (examples: oxygen used to help breathing)
	Borderline with/ MD or drugs
	Non Medical Device  

	28.
	Gingival gel
	Indicated use for gingivitis, gingival bleeding and other inflammatory and traumatic conditions where the gingival mucosa requires increase concentration of hyaluronic acid.
	Borderline with/ MD or drugs
	Medical Device 
Depends on the mode of action of the active ingredient. If it is physical mode of action then it is a Medical Device . But if it is pharmacological action, the it is classified as medicine.

	29.
	Contrast Media
	Used in radiodiagnostic procedures
	Medical Device  Borderline with Medicine
	Non Medical Device  

	30.
	Microscope
	Used in a variety of fields to view objects or organisms that are not visible with the naked eye. These devices are useful in biology, clinical and medicine to study or analyze specimens, such as bacteria, cells and body tissues
	Medical Device  or non Medical Device 
	Non Medical Device , if General Laboratory Purposes

Medical Device , if use for IVD, IVF

	31.
	Centrifuge
	A Laboratory device that is used for the separation of fluids, gas or liquid, based on density. Separation is achieved by spinning vessel containing material at high speed; the centrifugal force pushes heavier materials to the outside of the vessel.
	Medical Device  or non Medical Device 
	Non Medical Device , if General Laboratory Purposes

Medical Device , if use for IVD, IVF

	32.
	Vaccine/pharmaceutical/medical refrigerator
	Design to store vaccines and other medical products at a stable temperature (prevent temperature fluctuations) to ensure an ideal cold temperature storage for sensitive items due to quality issues and not degrade.
	Medical Device  or non Medical Device 
	Non Medical Device 
Medical Device  if used to refrigerate other Medical Device  (blood product)

	33.
	Control Reagent
	a reagent blanks that routinely processed with samples. Laboratories include reagent blanks in the extraction, quantitation, and amplification processes to aid in monitoring potential contamination of reagents and/or supplies
	Medical Device  or non Medical Device 
	In Vitro Medical Device 

	34.
	Calibrator Reagent
	Used in all instruments of measurement to establish a reference point that assists instruments to produce accurate results
	Medical Device  or non Medical Device 
	In Vitro Medical Device 

	35.
	General purpose reagent such as washing solution, detergent solution, buffer solution
	GPRs based on 21 CFR 864.4010(a) is a chemical reagent that has general laboratory application, that is used to collect, prepare and examine specimens from the human body for diagnostic purposes, and that is not labeled or otherwise intended for a specific diagnostic application
	Medical Device  or non Medical Device 
	Non Medical Device 

	36.
	Plaster with povidone iodine
	A wound dressing with povidone iodine is intended for use as a primary dressing for wounds. This is a dressing type where the povidone iodine additive is permanently bound to the textile substrate. The povidone iodine has antimicrobial activity.
	Borderline with/ MD or non MD
	Medical Device 

	37.
	Blood Bank Refrigerator
	For channeling and storing liquids for infusion
	Medical Device  or non Medical Device 
	Medical Device 

	38.
	Hemodialysis Dialysate
	Intended use for the exchange of solutes with blood through a semi-permeable membrane in the dialyser or a haemodialysis system
	Medical Device  or non Medical Device 
	Medical Device 

	39.
	Allograft
	a graft of tissue between individuals of the same species but of disparate genotype; types of donors   are cadaveric, livingrelated, and living unrelated
	Medical Device  or Biologics
	Non Medical Device 


	40.
	Aqueous cream product
	As an emollient cream with moisturizing properties to promote healing and relied to the symptoms of skin dryness, impaired barrier function, skin problems/ diseases.
	OTC Drug
	Depends on the mode of action of the active ingredient. If it is physical mode of action then it is a Medical Device . But if it is pharmacological action, the it is classified as medicine.

	41.
	Blood bag containing anticoagulant/ preservation agent
	To collect and preserve blood and its components (for use with cytapheresis device only)

Note: not for direct intravenous infusion.
	Medical Device  borderline with medicine (Device-drug combination product)
	Medical Device 

	42.
	Catheter lock/ Flush solutions (eg: heparinised saline, sodium citrate solution)
	As an anticoagulant for use as a catheter lock/ flush solution for flushing off catheters and cannulas to maintain catheter/ cannula patency and to prevent coagulation of blood or infection in the catheter.

Note: It is not indicated for therapeutic use, contraindicated for direct systemic administration
	Medical Device 
	Medical Device 

	43.
	Collagen Haemostatic agents (fibrillary or soft, pliable pad/ sponge or loose fibres)
	A sterile, bioabsorbable device derived from animal collagen (eg: bovine or porcine collagen) designed to produce a rapid haemostasis through platelet activation/ aggregation (which initiates the haemostatic cascade leading to a fibrin clot) during a surgical procedure. It is applied directly to the wound where it remains to be absorbed by the specific anatomy/application and does not contain an antimicrobial agent
	Medical Device 
	Medical Device 

	44.
	Peritoneal dialysis dialysate
	It is used for the exchange of solutes across the peritoneum of the patient (in this case, used as a semi-permeable membrane)
	Drug
	Non Medical Device 

	45.
	Haemodialysis dialysate
	It is used for the exchange of solutes with blood through a semi-permeable membrane in the dialyser of a haemodialysis system.
	Medical Device 
	Medical Device 

	46.
	Drug-eluting beads
	It is an embolic agent which is intended to be loaded with a chemotherapy agent.

Eg: Doxorubicin for the purpose of treatment of malignant hypervascularised tumors by embolization of vesses and occlusion of blood flow supplying malignant hypervascularised tumours and as a secondary action, delivers/elutes a local, controlled sustained dose of the chemotherapy agent directly to the tumours.
	Medical Device  borderline with medicine.

a. if the beads and drug are packaged and sold together, it will classified as Drug-device combination.

b. If the beads are sold separately from the drug, it will classified as Medical Device 
	Non Medical Device 
*Medical Device  if the beads are sold separately

	47.
	Drug-eluting Stents (DES)
	For use in angioplasty or coronary stenting procedures.
	Medical Device  borderline with medicine
	Medical Device 

	48.
	Drug-delivery products 

(eg: insulin prefilled pen/ syringes, asthma inhalers)
	To administer pharmacologically active substance
	Medical Device  Borderline with Medicine (Drug-device combination product)
	Non Medical Device 
*Medical Device  if the devices are sold separately

	49.
	Enteral Feeding Kit (containing Iodine Pack drug)
	A collection of sterile devices that includes tubing and other materials intended to administer nutrient liquids directly into the stomach, duodenum or jejunum of a patient by means of gravity or enteral pump.
	Medical Device  Borderline with Medicine (Device-drug combination product)
	Medical Device 


	50.
	Aqueous/ vitreous humour replacement medium
	It is used to assist in performing ophthalmic surgery (eg: to maintain the shape of the eyeball during the intervention, preserve tissue integrity, protect from surgical trauma, or to function as a tampnade during retinal reattachment
	Medical Device 
	Medical Device 

	51.
	General purpose surgical or barrier drapes
	A sterile protective covering made of natural or synthetic materials.

To isolate a site of surgical incision or a surgical field from contamination (eg: microbial substance) in various clinical settings.

The device may also be used to protect a patient from heat/flame during a surgical procedure.
	Medical Device  Borderline with Medicine (Device-drug combination product)
	Medical Device 

	52.
	General-body orifice lubricant
	Lubricant intended to facilitate entry of a diagnostic or therapeutic device into a body orifice by reducing friction between the device and the body;

Lubricant during catherisation, probing, endoscopy, changing fistula catheters, intubation and prevention of iatrogenic injuries to the rectum and colon.

Eg: ancillary local anaesthetic: lidocaine
	Medical Device  Borderline with Medicine (Device-drug combination product)
	Medical Device 

	53.
	Heat pad/ cooling pad
	To relief aches and pains
	Medical Device 
	Medical Device 

	54.
	In vivo diagnostic agents
	For diagnostic purpose

Eg: 

- X-ray/ MRI contrast media

- NMR enhancing agents

- Opthalmic diagnostic agents

- Carrier solutions to stabilize microbubbles for ultrasound imaging

- Radiopharmaceuticals for diagnostic use
	Drug
	Non Medical Device 

	
	
	As diagnostic test kit consist of drug and analyser, example Urea Breath Test
	Medical Device  Borderline with Medicine (Drug-device combination product)
	The Analyser Instrument is Medical Device 
The Urea Tablet is Drug

	
	
	As diagnostic analyser only (without drug)
	Medical Device 
	Medical Device 

	55.
	Medical gases
	a. To be used in anaesthesia and inhalation therapy, including their primary containers

b. For in-vivo diagnostic purposes including lung function tests.
	Drug
	Non Medical Device 

	56.
	Medicinal patch
	To relieve fatigue, body aches, joint pains; to regulate hormone imbalance
	Drug
	Non Medical Device 

	57.
	Preparations (including

swabs/ wipes) containing

antiseptics/ disinfectants/

antimicrobial substances
	Intended for the disinfection of Medical Device s.

(eg. alcohol, chlorhexidine, iodine, cetrimide)
	Medical Device 
	Medical Device 
16th AMDTC Meeting agreed to remove “antiseptics”. 

	58.
	Peeling/ Exfoliator product
	To improve skin texture due to unaesthetic skin appearance caused by pigmentation, post acne scars, photo damage, etc.

(eg. Products containing glycolic acid and salicylic acid)

Note:

The ingredient and intended use should comply with the Guidelines for Control of Cosmetic Products in Malaysia.
	Cosmetic (Drug)
	Non Medical Device 

	59.
	Personal intimate hygiene
	a. For female/male intimate hygiene

Note:

The product should be rinsed off.

b. For symptomatic relief of vaginal irritation/infections by changing the vaginal pH.
	a. Cosmetic (Drug)

b. Drug
	Non Medical Device 


	60.
	Soft tissue filler/ Dermal filler
	To correct cutaneous contour deformities of the skin (e.g., moderate to severe facial wrinkles and folds such as nasolabial folds, scars), particularly in cases of aging or degenerative lesions
	Medical Device  Borderline with Medicine (Device drug combination product)
	Medical Device 

	61.
	Synthetic fluid tissue reconstructive material
	As a submucosal implant in the urinary tract for urinary incontinence or vesicoureteral reflux.

It may also be injected into the vocal cords to treat the effects of paralysis, atrophy, or scarring. After application, this device cannot be reused.
	Medical Device  Borderline with Medicine (Device drug combination product)
	Medical Device 

	62.
	Wart Products

(eg. pen applicator containing a caustic agent, cyryogenic kit with refrigerant)
	a. Containing a caustic agent eg. trichloroacetic acid (TCA) that destroys warts by chemical coagulation of proteins.

b. Cryotherapy which destroys warts by freezing them using a very cold substance eg: liquid nitrogen or refrigerant made from dimethyl ether and propane.
	a. Drug

b. Medical Device 
	a. Non Medical Device 
b. Medical Device 

	63.
	Comprising a matrix (eg. dressing, gauze, swabstick, plaster, sponge)
	a. To administer a medicinal substance to the wound eg: antimicrobial/ antiseptic agent for the purpose of controlling infection.

b. To provide a protective layer/barrier to the wound and prevent microbial penetration and create healing environment. It may incorporate an ancillary medicinal substance eg: antimicrobial/ antiseptic agent.
	a. Drug

b. Medical Device 
	a. Non Medical Device 
b. Medical Device 

	64.
	Topical preparation for application to a skin wound

(e.g., abrasion, laceration, cut, ulcer)
	To facilitate local haemostasis. It is available in various forms (e.g., gel, spray, powder, ointment, plaster/gauze pad) that can be applied directly to the wound where it forms a seal of transparent layer.
	Medical Device 
	Medical Device 

	65.
	Deep cavity wounds dressing for application to a surgical wound
	To use as the wound covering material for deep body cavity to reduce the adhesion of surrounding tissues by applying to the surgical area
	Medical Device 
	Medical Device 

	66.
	Intravascular catheter securement device containing antimicrobial/antiseptic agent (e.g. chlorohexidine gluconate, CHG)
	An intravascular catheter securement device is a device with an adhesive backing that is placed over a needle or catheter and is used to keep the hub of the needle or the catheter flat and securely anchored to the skin. The antimicrobial agent provides ancillary antimicrobial activity to reduce skin colonization and catheter colonization, supress regrowth of microorganism’s, and reduce catheter-related bloodstream infections (CRBSI) in patients with central venous or arterial catheters.
	Medical Device  Borderline with Medicine (Device-drug combination product)
	Medical Device 

	67.
	Surgical apparel
	Example: shoe cover, surgical gown/apron, cap
	Medical Device 
	Medical Device 

	68.
	Equipment cover and Equipment sleeves for operating theatre
	To cover equipment for operating theatre
	Not Medical Device 
	Medical Device 

	69.
	Trolley
	A trolley to organize and storing medications and transporting to patient easily/ to transport medicine, instruments and disposable materials/ transport utility/ transport soiled linen/clothing to the laundry
	Not Medical Device 
	Non Medical Device 

	70.
	Examination stool
	Stool for doctors to sit for examination purpose.
	Not Medical Device 
	Non Medical Device 

	71.
	Sleeper chair
	A chair framework constructed from tublar steel and with pvc upholstery. Can be converted from a chair to a sleeper unit for sitting and temporary sleeping for hospital visitors.
	Not Medical Device 
	Non Medical Device 

	72.
	Urinary bottle holder
	Stainless steel framework to hold urinal bottle
	Not Medical Device 
	Non Medical Device 

	73.
	Dental bib
	Dental cover for patient’s body
	Not Medical Device 
	Non Medical Device 

	74.
	Identity band, markers, padlock security tags, container label
	For tagging purpose, time markers is used to follow the precise part of food through the entire GI tract and specifically in the bowel, whist, not influencing the performance of normal gastric and intestinal function, for surgical instrument container
	Not Medical Device 
	Non Medical Device 

	75.
	Cart
	Example: Medical cart, laundry cart, linen cart
	Not Medical Device 
	Non Medical Device 

	76.
	Dental aprons
	
	Not Medical Device 
	Non Medical Device 

	77.
	Reclining chairs
	Chairs for visitors and patient to rest
	Not Medical Device 
	Non Medical Device 

	78.
	Plaster with Silver
	It is used for sterilizing & protecting contamination of a wound site such as pressure ulcers, leg ulcer, diabetic ulcer, traumatic wounds, lacerations, donor and graft sites, burns, post-operative wound, and venous stasis ulcer.
	Drug
	Medical Device  

	79.
	Test equipment’s and refills
	The equipment’s and refills are for QC purpose
	Medical Device 
	Non Medical Device 

	80.
	Sterilization Indicator
	Sterilization indicator as accessories

Example: steam sterilization, concentration of peracetic acid/ peroxide
	Medical Device 
	Medical Device  (Accessory)

	81.
	IVD Purpose Benchtop incubator
	For cell culture purpose, only for IVD purpose
	
	Medical Device 

	82.
	CPAP Mask Wipes
	CPAP, or continuous positive airway pressure, is a treatment that uses mild air pressure to keep the airways open. CPAP typically is used by people who have breathing problems, such as sleep apnea. CPAP also may be used to treat preterm infants whose lungs have not fully developed. The Wipes is use to disinfect the CPAP mask.
	Medical Device 
	Medical Device  if it is use for disinfecting

	83.
	Ward screening
	To provide temporary or permanent privacy solution
	Not Medical Device 
	Non Medical Device 

	84.
	Lice comb
	To remove lice and lice eggs from human hair
	Medical Device 
	Non Medical Device 

	85.
	Invasive blood pressure accessories
	The invasive blood pressure adapter cables are intended to be used for connecting one or two invasive blood pressure transducers to get medical systems information technologies and datex-ohmeda monitoring equipment requiring rectangular 11-pin connector cable.


	Medical Device 
	Medical Device 

	86.
	Mayo table
	Surgical mayo instrument stand
	Not Medical Device  
	Non Medical Device 

	87.
	Lichensed Adult Syrup
	Lichensed adult syrup acts as a barrier by creating a protective film that soothes a cough and protect the airway.
	Drug
	Non Medical Device 

	88.
	Monitor cable
	To connect the disposable pressure transducer to the patient monitor
	Medical Device  
	Medical Device  System Accessory

	89.
	Cold laser therapy
	Treatment of pain management, accelerate healing
	Medical Device  
	Medical Device 

	6th AMDTC

	90.
	Diapers
	
	Medical Device   or non-Medical Device   
	Non-Medical Device  

	91.
	Milk bottle and Nipple
	
	Medical Device   or Food
	Non-Medical Device  

	92.
	Breastmilk storage and freezing bags
	
	Medical Device   or Food
	Non-Medical Device  

	93.
	Pacifiers
	
	Medical Device  or non-Medical Device  
	Non-Medical Device 

	94.
	Disposable breast pad
	
	Medical Device  or non-Medical Device  
	Non-Medical Device 

	95.
	Nasal Aspirator
	
	Medical Device  or non-Medical Device  
	Medical Device 

	96.
	Urine bags and vomit in children or adults
	Incorporation into a deodorant gel can be used at home or outdoors, disposable.
	Medical Device  or non-Medical Device  
	Non-Medical Device 

	97.
	Glasses frames
	Glasses frames without lenses.
	Medical Device  or non-Medical Device  
	Non-Medical Device 

	
	
	Glasses frames that come with lens but no power.
	Medical Device  or non-Medical Device  
	Non-Medical Device 

	98.
	Sunglasses
	Sunglasses that come with lens but no power.
	Medical Device  or non-Medical Device  
	Non-Medical Device 

	99.
	Spectacle lens has no power
	
	Medical Device  or non-Medical Device  
	Non-Medical Device 

	100.
	Semi-finished spectacle lens blank
	
	Medical Device  or non-Medical Device  
	Non-Medical Device 

	101.
	Powered contact lens
	Soft and hard lens
	Medical Device 
	Medical Device  


	102.
	Disinfecting Lens care solution
	​​
	Medical Device  or drug
	Medical Device  



	103.
	Ophthalmic Viscosurgical Device (OVD)
	It contains substances with viscosity properties such as Sodium hyaluronate, Chondroitin sulfate, Hydroxypropylmethylcellulose (HPMC)
	Medical Device  or drug
	Medical Device  

	104.
	Keratome blade
	For eye surgery
	Medical Device 
	Medical Device 

	105.
	The eye surgery simulator
	For education or lesson
	Medical Device  or non-Medical Device  
	Non-Medical Device 

	106.
	Colour blindness tests
	For diagnostic colour blindness
	Medical Device  or non-Medical Device  
	Medical Device 

	107.
	Optical lens manufacturing machine
	
	Medical Device  or non-Medical Device  
	Non-Medical Device 

	108.
	Lensmeter
	Use by employee inspection department template at the lens manufacturing plant.
	Medical Device  or non-Medical Device  
	Non-Medical Device 

	109.
	Autopsy table, Autopsy saws, Morgue refrigerator, Body bag
	
	Medical Device  or non-Medical Device  
	Non- Medical Device 

	110.
	Sterile blood bag
	Sterile empty blood bag
	Medical Device  or Drug
	Medical Device 

	111.

	Collection blood tube
	Empty collection blood tube
	Medical Device 
	Medical Device 

	
	
	Collection blood tube with anticoagulation
	Medical Device  or Drug
	Medical Device 

	112.
	Plasma Freezer, Blood product freezer
	
	Medical Device  or non-Medical Device 
	Medical Device 

	113.
	Haematocrit Centrifuge
	Spin the specimen from the human and obtain information from medical examinations or diagnostic purposes.
	Medical Device  or non-Medical Device 
	Medical Device 

	
	Centrifuge for IVD
	Spin the specimen from the human body
	Medical Device  or non-Medical Device 
	Medical Device 

	114.
	Platelet Additive Solution
	To replace the plasma portion of the preparation and storage buffy coat-derived platelet concentrate orapheresis platelet units To feed and maintain the condition of platelets.
	Medical Device  or Drug
	Medical Device 
16th AMDTC Meeting agreed to remove this item. 

	115.
	White blood cell type T-lymphocyte counting reagent
	Used in conjunction with hematologic analytical automation
	Medical Device  or non-Medical Device 
	Medical Device 

	116.
	Control Solution
	Control substances used to verify the accuracy of Medical Devices.
	Medical Device  or non-Medical Device 
	Medical Device  if it is use for IVD

	117.
	Pedometer
	Walking motion detector or step counter and tell the amount of energy and fat that is metabolized or tell the distance and time between walking or calorie processing used to plan a diet.
	Medical Device  or non-Medical Device 
	Non-Medical Device 

	118.
	Wristwatch
	Measure heart rate, steps, body fat, Breathing rate/respiratory rate, mood, stress, deep sleep while exercising or daily life and not for medical purposes
	Medical Device  or non-Medical Device 
	Non-Medical Device 

	
	
	Measure oxygen saturation, blood glucose, blood pressure, body temperature, ECG sensor
	Medical Device  or non-Medical Device 
	Medical Device 

	119.
	Treadmill
	For exercise/In fitness
	Medical Device  or non-Medical Device 
	Non-Medical Device 

	
	
	For diagnostic cardiovascular
	Medical Device  or non-Medical Device 
	Medical Device 

	120.
	Bone cement
	
	Medical Device  or Drug
	Medical Device 

	121.
	Antibiotic Bone cement

	Antibiotics such as Gentamicins
	Medical Device  or Drug
	Medical Device 

	122.
	Vacuum Splint Kit
	Prevents bones moving

	Medical Device 
	Medical Device 

	123.
	Back support


	Prevent back pain and bones move if bent or lift heavy or wrong for medical purposes.
	Medical Device  or non-Medical Device 
	Medical Device 
16th AMDTC Meeting agreed to replace this item with “orthosis support”.

	7th AMDTC

	124.
	X-ray and Gamma Radiation Personal Dosimeter
	For safety in work areas with radiation, such as radiation laboratory.
	MD or non-MD
	Non-Medical Device



	125.
	X-ray Monitor
	Both static animated images in color and monochrome images from various medical instruments such as CT Scan, MRI, Digital Mammogram, Ultrasound, PET Scan. Technical radiation is a visual validator. Radiologist is the reader with the interpretation of the image. And the doctor is the viewer. With the results of the examination. Analysis to plan for further treatment.
	MD or non-MD
	Medical Device



	126.
	X-ray protective clothing such as  Gown, Aprons, Collar, Glasses, Gloves
	Use of industrial radiation protection
	MD or non-MD
	Non-Medical Device


	127.
	Shield protects radio waves from external environment into the diagnostic room.
	To prevent radio interference from MRI, Used to provide the right place or environment for the MRI machine.
	MD or non-MD
	Non-Medical Device



	128.
	Pastille or spray for people with symptoms hyposalivation or xerostomia used in conjunction with treatment mucositis due to radiation or high doses of chemotherapy.
	To help moisturize, moisturize and cleanse the mouth for people with dry mouth. The mechanism of physical action.
	MD or Drug
	Medical Device



	129.
	Dressing to reduce inflammation such as dry, itching, flaking, peeling irritated, pain , redness, heat from radiation therapy or first and second degree burns
	The cover helps to keep the wound healed faster and mechanism is physical action.
	MD or Cosmetics
	Medical Device



	130.
	Metal Detector, Seal Inspection, Weigh Inspection
	Used in various industries such as medicine, food, cosmetics, medical equipment, automotive, plastics, etc.
	MD or Non-MD
	Non-Medical Device



	131.
	Facial imaging equipment
	Device to check the condition such as oily skin  or dry skin
	MD or Non-MD
	Non-Medical Device



	132.
	Skin detector equipment
	To see the size, pores, skin color, wrinkles, dryness, dark spots, and to be used as a guide to use cosmetics for cosmetic customers.
	MD or Non-MD
	Non-Medical Device



	133.
	Machine or device that has a brush or sponge (puff) that can rotate or shake
	Use cleaning polish or for spread out of cream
	MD or Non-MD
	Non-Medical Device



	134.
	Machine or equipment sprayed on water or mineral water 
	To keep the face moist
	MD or Non-MD
	Non-Medical Device



	135.
	Products using BIA (Bioelectrical Impedance Analysis) working principle
	Body measuring instrument It is intended for measuring and analyzing body composition such as body weight, fat mass, muscle mass Total body weight, body mass index, body fat percentage, body fat The body's metabolic rate, waist to hip ratio. Body age estimate Compare the age of the body Exercise plan. Without a medical purpose or diagnosis.
	MD or Non-MD
	Non-Medical Device



	136.
	Motion Capture System
	To record human movements, animals and inanimate objects. To study movement 3D animation simulation used in the entertainment industry, such as movies, movies, games, sports industry.
	MD or Non-MD
	Non-Medical Device



	137.
	Baraku, Electronic Baraku, Electronic Cigarette
	
	MD or Non-MD
	Non-Medical Device



	138.
	Medical mask, Surgical Mask, Surgical respirators
	For medical use
	MD or Non-MD
	Medical Device

	139.
	Disposable mask for Industry, Firefighter mask, Disposable dust mask
	Not for medical use
	MD or Non-MD
	Non-Medical Device



	140.
	Nasal Strips
	Expansion of respiratory tract makes breathing easier to reduce snoring
	MD or Non-MD
	Medical Device

	141.
	wrist strap, ankle strap
	To specify the details of the patient or infant, such as surname, date of treatment, etc.
	MD or Non-MD
	Non-Medical Device



	142.
	Insoles or contain activated charcoal
	Put on a shoe to keep your feet comfortable or reduce unwanted odors.
	MD or Non-MD
	Non-Medical Device



	143.
	Heel liners
	Prevent bite shoes or to tighten. Without a purpose for medical use.
	MD or Non-MD
	Non-Medical Device



	144.
	Foot cushion
	Supports shock and friction caused by foot pressure. Without a purpose for medical use.
	MD or Non-MD
	Non-Medical Device



	145.
	Tools or equipment for measuring the size and shape of individual feet
	It is used to cut shoe or shoe to fit each foot
	MD or Non-MD
	Non-Medical Device



	146.
	Blade for shaving, Epilator hair removal, Temporary hair removal, Shaver, Machine or wax melting pot used to remove hairs
	
	MD or Non-MD
	Non-Medical Device



	147.
	Suture Needle for wound closer
	
	MD or Non-MD
	Medical Device

	148.
	Suture Needle for lifting
	For lifting face, chin, body, etc.
	MD or Non-MD
	Medical Device

	149.
	Tattoo Machine
	For cosmetics or for belief or religion
	MD or Non-MD
	Non-Medical Device

	150.
	Tattoo Needles, Tattoo Ink
	For cosmetics or for belief or religion
	MD or Non-MD
	Non-Medical Device

	151.
	Piercing Needles, Clamping device, grab jewelry
	Use to pierce the ear, nose, mouth, tongue, navel, nipple or other areas to wear jewelry. And for beauty or for belief or religion
	MD or Non-MD
	Non-Medical Device

	152.
	Ambulance, Emergency Vehicle, Mobile Medical Check-up Vehicle,       Mobile X-Ray Vehicle,        Mobile Dental Vehicle,     Mobile Laboratory Vehicle
	
	MD or Non-MD
	Non-Medical Device

	153.
	The sheet or device records the time and temperature of the environment surrounding the product while it is being shipped.
	For information on how to keep the product from damage from high or low temperatures.
	MD or Non-MD
	Non-Medical Device

	154.
	Massage Bath Tub
	Used in entertainment type massage parlor. To entertain or relax the user.
	MD or Non-MD
	Non-Medical Device

	155.
	Couch
	For sleep to support the face. Body massage or for beauty
	MD or Non-MD
	Non-Medical Device

	156.
	Swimming Pool Spa and Outdoor Spa
	The pump system. Heating system water pressure system, ozone system, fountain system, music and LED lights are intended for soaking water to massage the body and help to relax. And can create atmosphere with music and light. Available at home, spa, clinic, hospital.
	MD or Non-MD
	Non-Medical Device

	157.
	Disposable urinal
	Urine bags are disposable. To use while traveling Stay overnight for general people, patient or disabled
	MD or Non-MD
	Non-Medical Device

	158.
	Thermometer in sauna
	Installed in sauna to measure temperature.
	MD or Non-MD
	Non-Medical Device

	159.
	Hygrometer in sauna
	Installed in Sauna Room to measure relative humidity.
	MD or Non-MD
	Non-Medical Device

	160.
	Motorized bed
	For sleeping at home adjustable height, up, adjust the head or leg lift.
	MD or Non-MD
	Non-Medical Device

	161.
	Electric scooter
	For driving to facilitate short trips. Used in places such as home, office, shopping mall, Exhibition Center Park
	MD or Non-MD
	Non-Medical Device

	162.
	Surgical warming blanket
	Used to warm patients during surgery.
	MD or Non-MD
	Medical Device

	163.
	Heated underblanket
	Use it on the mattress and cover sheets to make it. Mattress on warm bed Used in cold weather. Without the purpose of medical use.
	MD or Non-MD
	Non-Medical Device

	164.
	Thermal overblanket
	Electric fields and magnetic fields within the product release heat to warm the body. Used in cold weather. Without the purpose of medical use.
	MD or Non-MD
	Non-Medical Device

	165.
	Temperature Boiler
	Connect with the mattress to adjust the temperature of the mattress to warm up. Used in cold weather.
	MD or Non-MD
	Non-Medical Device

	166.
	Hot pad or Hot bag
	Put on clothes Under the socks to warm the body. Suitable for those who need warmth in cold weather. Or those who play winter sports without the purpose of medical applications.
	MD or Non-MD
	Non-Medical Device

	167.
	Application - Body fat
	Not for medical purpose
	MD or Non-MD
	Non-Medical Device

	168.
	Models, Simulators and Medical Skills Training Products
	The product is not actually available in patients or those with abnormalities.
	MD or Non-MD
	Non-Medical Device

	169.
	Microscope to see the tissue of the human body
	Especially to diagnose diseases or injuries of humans.
	MD or Non-MD
	Medical Device

	170.
	The microscope to see size, quantity and visualizes the internal structure of a small object.
	For study or research such as cell, tissue, wood, rock, mineral, animal secretion, foam, snow, sample metal from the culture plate, plant, insect, small animal.
	MD or Non-MD
	Non-Medical Device

	171.
	Product cover on microscope to prevent moisture.
	To prevent moisture that causes fungus in the microscope.
	MD or Non-MD
	Non-Medical Device

	172.
	Electronic Tables
	Electrical equipment or items use in hospitals, clinics or other places.
	MD or Non-MD
	Non-Medical Device

	173.
	Medicine Trolley, UD Dose Cart
	To distribute medicine to each patient.
	MD or Non-MD
	Non-Medical Device

	174.
	Mobile Computer Cart
	Can be moved to other places used in hospitals, clinics, offices.
	MD or Non-MD
	Non-Medical Device

	175.
	Device for patient or central pipeline system.
	Flowmeter, Humidifier, Nebulizer, Suction Tube, Nasogastric Tube, Suction For Surgical Room
	MD or Non-MD
	Medical Device

	8th AMDTC

	176.
	Anti-virus patch
	Intended to be attached to cellphones, IDs, bags, etc.  The device releases chlorine dioxide that eliminates respiratory viruses.  Hence, preventing virus transmission
	MD or non-MD
	Non-Medical Device

	177.
	Rectal Cream/Suppository
	Intended to soothes and moisturize the perianal area, relieves and rejuvinates common signs of hemorrhoids, and relieve burning sensation or itchiness in the problematic area. NOTE: No drug components only emollient, emulsifying and film forming materials such as stearic acid, hipericum oil and beeswax
	MD or non-MD
	Medical Device

	178.
	Pimple patch
	With hydrocolloid technology to conceal the pimple, helps avoid bacterial contact, helps to dry out and flatten the surface level of the  acne NOTE:  components are, sodium carboxycellulose, elastomer, dioctyl adipate, hydrogenated petroleum resin
	MD or non-MD
	Medical Device

	179.
	Diamond impregnated stone for trimming, zirconia polishers, polisher disc, CAD/CAM miling burss, milling unit, ultra vaccum, grind control basic unit
	Intended to be used in the dental laboratory by a laboratory technician for the manufacture of dentures, crowns, and bridges
	MD or non-MD
	Non-Medical Device

	180.
	Oxygen recreational
	Use supplemental recreational oxygen to help recover from stress in the body due to exertion. The oxygen seems to help recovery, give more stamina, and provide an energy boost when needed.
	MD or non-MD
	Non-Medical Device

	9th AMDTC

	181.
	Medical drape
	To be put on treatment table or bed during operation procedure and to prevent transfer of microorganisms, fluids and particles to the patients
	MD or non-MD
	Medical Device

	182.
	Mattress
	Combine a specific foam density with a nominal hardness that optimises pressure reduction and improves circulation and quality of sleep.
	Medical Device or Non-Medical Device


	Non-Medical Device

	183.
	Therapressure brush
	To decrease the state of arousal (calm), normalize reponses to touch, help the child relax and become more aware of his/her body to improve attention and purposeful movements.
	Medical Device or Non-Medical Device


	Non-Medical Device



	184.
	Portable oxygen spray
	Recreational portable canned oxygen can help improve stamina, increase endurance, and help muscles recover faster from soreness and fatigue.
	Medical Device or Non-Medical Device


	Non-Medical Device



	185.
	Air purifier
	Air purifier that traps and eliminates bacteria and viruses found in the air, effectively reducing the risk of secondary infection and airborne transmission. For home use.
	Medical Device or Non-Medical Device


	Non-Medical Device



	186.
	Smartwatch
	Measure skin temperature (only area that contact with smartwatch). Not intend to use for diagnostic or medical purpose.
	Medical Device or Non-Medical Device
	Non-Medical Device

	10th AMDTC

	187.
	Strip to check saliva acidity
	To measure saliva acidity for diagnosis of teeth caries
	Medical Device or Non-Medical Device
	Medical Device

	188.
	Microtube
	Empty specimen container for centrifuge and PCR 

	Medical Device or Non-Medical Device
	Medical Device for IVD use

Non-Medical Device for laboratory use

	189.
	Solution contains Trichloroacetic Acid (TCA)
	Face skin peeling, removal of death cell
	Medical Device or Non-Medical Device
	Non-Medical Device

	190.
	Solution contains lactic acid, glycolic acid, and salicylic acid
	Face skin peeling, removal of death cell
	Medical Device or Non-Medical Device
	Non-Medical Device

	191.
	3D Printer
	To produce orthodontic models. For dental laboratory use. 
	Medical Device or Non-Medical Device
	Non-Medical Device

	
	
	To fabricate Custom-Made Bones from non-degradable materials such as polymer or titanium
	Medical Device or Non-Medical Device
	Non-Medical Device

	
	
	3D bioprinter - fabricate a dermal patch graft to treat scarred lesions from biomaterial such as patient's autologous tissues and cells
	Medical Device or Non-Medical Device
	Non-Medical Device

	192.
	Allograft bone
	Intended for replacement of damaged bone/ transplantation

from: 

6.1 morselized frozen bone 

6.2 frozen bone

6.3 freeze-dried bone
	Medical Device or Non-Medical Device
	Non-Medical Device

	193.
	Device for production of acidic dialysis concentrate solution
	Device contains the hemodialysis powder that is mixed with RO water and transferred to the storage tanks ready to be fed directly to the hemodialysis machines for the patient.
	Medical Device or Non-Medical Device
	Medical Device

	194.
	Equipment that releases Negative Ions
	Intended for treatment of menstrual cramps, balance hormones, treatment of glaucoma.
	Medical Device or Non-Medical Device
	Medical Device

	11th AMDTC 

	195.
	Heating & cooling massager
	39°C heating mode helps to improve skin circulation and cleanse skin residues deeply by opening pores.

36°C booster mode helps skin absorption of cosmetic with micro vibration.

13°C cooling mode makes your skin lower temperature immediately and keeps the optimal skin temperature.

5°C tightening mode with cold massage. Shrinks pores and makes skin elastic.
	Medical Device or Non-Medical Device


	Non-Medical Device

	196.
	Antibacterial Surface Coating (Not for Medical Device coating)
	Disinfection coating for surfaces, the material is silver or copper.
	Medical Device or Non-Medical Device


	Non-Medical Device

	197.
	Wound Antiseptics
	Topical antiseptic wound cleansers aim to prevent infection due to wounds of the skin.
	Medical Device or Non-Medical Device


	Non-Medical Device

	198.
	Air disinfectant (hospital use)
	Air cleaning liquid (disinfectant) which can be used by spraying it into the air so that it forms smoke or mist.
	Medical Device or Non-Medical Device


	Non-Medical Device

	12th AMDTC 

	199.
	Software
	1.1 Standalone software that intended to display, receive, collect, and store health data (such as blood pressure, blood glucose) with no processing or analysis for wellness.
	Medical Device or Non-Medical Device
	Non-Medical Device

	
	
	1.2 Standalone software that intended to collect patient medical data by connected directly with medical device and transmit the patient medical data to a doctor/physician for further consideration without processing and analyzing the data
	Medical Device or Non-Medical Device
	Non-Medical Device

	
	
	1.3  Standalone software that intended to collect patient medical data by not connect to medical device (e.g. filling the information by user) and transmit the patient medical data to a doctor/physician for further consideration without processing and analyzing the data
	Medical Device or Non-Medical Device
	Non-Medical Device

	200.
	Electronic cigarette 
	An electronic cigarette is an electronic device that simulates tobacco smoking
	Medical Device or Non-Medical Device
	Non-Medical Device

	201.
	Sex toy 
	An object or device used for sexual stimulation or to enhance sexual pleasure.
	Medical Device or Non-Medical Device
	Non-Medical Device

	202.
	Telecare Alarm System 
	Intended to detect fall from hospital bed, general bed and alert to Central Monitoring Station e.g. hospital, elderly care, owner company 
	Medical Device or Non-Medical Device
	Non-Medical Device

	203.
	Fall Detection Devices
	Wearable sensor devices (e.g. smartwatch, band) or ambient sensor device intended for fall detection.
	Medical Device or Non-Medical Device
	Non-Medical Device

	204.
	Phantom for Quality Assurance of Radiological Product
	Intended as radiotherapy quality assurance for determination radiation dose from Medical linear accelerator (LINAC), Intended for calibration of the product
	Medical Device or Non-Medical Device
	Non-Medical Device

	13th AMDTC

	205.
	Nasal Spray 
	Intended to be used  for nasal cleaning containing deep sea salt water.
	Medical Device or Non-Medical Device
	Medical Device

	206.
	Wound cover gel/cream
	Containing hyaluronic acid or sodium hyaluronate and has the following effects: Covering acute and chronic wounds (first and second degree burns, vascular and metabolic ulcers and pressure ulcers) and provide a moist wound environment free of bacteria.
	Medical Device or Non-Medical Device
	Medical Device

	14th AMDTC

	207.
	Reverse osmosis equipment
	used in hemodialysis water treatment systems: produces water to dilute concentrates used in hemodialysis
	Medical Device or Non-Medical Device
	Medical Device

	208.
	Sprays and creams help prolong sexual intercourse
	The product does not contain pharmaceutical ingredients, containing Oat Amino Acids, Aloe Vera, Alfalfa Extract, Borago Officinalis Extract, Chamomile Extract, Clove Essential Oil
	Medical Device or Non-Medical Device
	Non-Medical Device

	209.
	Dermal implant
	For Treatment  of facial contour deformities(such as asymmetry treatment) in case of tissue loss by accident or  disease, to add volume into facial dermis for treatment of facial  lipoatrophy. contains PDO polydioxanone. 
	Medical Device or Non-Medical Device
	Medical Device

	210.
	Nitric oxide Nasal spray
	-an active barrier that shields and protects from airborne viruses found often in circulated air in open spaces, planes, trains, large shopping centre and public transportation
	Medical Device or Non-Medical Device
	Medical Device

	
	
	-Use as a protective mechanical barrier against viruses within the nasal cavity
	
	

	15th AMDTC 

	211.
	Disinfectants for dental equipment
	Used in dentistry to disinfect dental instruments containing the following ingredients: chlorhexidine gluconate, iodine, formaldehyde, sodium hypochloride
	Medical Device or Non-Medical Device
	Medical Device

	212.
	Vaginal gel
	Used as an aid in the prevention and treatment of vaginal atrophy. The gel moisturizes and lubricates the vaginal mucosa, reducing symptoms of dryness, burning, itching and discomfort during sexual intercourse. Can reduce symptoms of vaginal atrophy, vaginal dryness and irritation at the time of use and long-term.


Mode of action: Maintains natural hydration of vaginal mucosa, relieves symptoms of dryness, burning, itching and discomfort during sexual intercourse.
	Medical Device or Non-Medical Device
	Medical Device

	213.
	Vaginal spray gel
	Help control of physiological conditions of the transformation zone of the cervico-vaginal mucosa to support prevention and recovery of cervical lesions caused by Human Papillomavirus.

Mode of action:

· Form a film protecting the cervico-vaginal mucosa from the deleterious action of external microbial agents, thanks to its capacity of bioadhesion.

Promote re-epithelialisation and regression of cervical lesions, thanks to the control of physiological conditions on the cervico-vaginal transformation zone by hydration.
	Medical Device or Non-Medical Device
	Medical Device if the mode of action is physical

	214.
	Fertilization Medium
	Medium for culture of embryos for In vitro Fertilization.
	Medical Device or Non-Medical Device
	Medical Device

	215.
	Radiopharmaceuticals Agents (Dimercaptosuccinic acid, Sodium Pertechnetate Tc -99m Injection, Diethylenetriaminepentaacetic acid, Methylenediphosphonic acid, Ethylene dicysteine, Sestamibi)
	Radiodiagnostic pharmaceutical products

Radiopharmaceuticals are agents used to diagnose certain medical problems
	Medical Device or Non-Medical Device
	Non-Medical Device

	216.
	Keratolytic agent
	The product group that exhibits its mechanism of action as a keratolytic agent, such as salicylic acid and trichloroacetic acid (TCA), is used for therapeutic purposes, including the treatment, relief, and management of conditions such as warts and nail fungus.
	Medical Device or Non-Medical Device
	Non-Medical Device

	217.
	Microneedle patch, dissolvable
	Dissolvable Microneedle Patch is
a microneedle patch designed for cosmetic applications, composed of cosmetic-grade ingredients that gradually dissolve within the skin. This allows beauty-enhancing components to penetrate the stratum corneum effectively. Each microneedle is 100 µm in length.
	Medical Device or Non-Medical Device
	Non-Medical Device

	218.
	Microneedle patch, non-dissolvable
	Non-Dissolvable Microneedle Patch is
a microneedle patch designed for cosmetic applications, made from biocompatible materials. It facilitates the enhanced absorption of skincare ingredients into the skin. Each microneedle is 100 µm in length.
	Medical Device or Non-Medical Device
	Non-Medical Device

	219.
	Endotoxin Assay Kit
	To determine the amount of endotoxin in WFI (Water for Injection) and injectable products manufactured by pharmaceutical factories.
	Medical Device or Non-Medical Device
	Non-Medical Device

	220.
	Drying Oven for laboratory equipment
	Sterilization or cleaning using dry air, used to sterilize glassware such as Erlenmeyer flasks, Petri dishes, test tubes, and other glass equipment.
	Medical Device or Non-Medical Device
	Non-Medical Device

	221.
	Patient restraint
	Device or technique that limits a patient's movement. Restraints are used to keep patients safe and prevent them from harming themselves or others. Specifically used for certain patient conditions, such as patients with mental disorders. It is usually made of fabric.
	Medical Device or Non-Medical Device
	Non-Medical Device

	222.
	Dental Sintering Furnace
	Creating various types of dental restorations, including crowns, bridges, and implants. The furnace can handle different types of ceramics and materials, making it an essential tool for any dental laboratory.
	Medical Device or Non-Medical Device
	Non-Medical Device

	223.
	Gas Flow Analyzer
	Measuring the output test of a ventilator; the device is only used during ventilator calibration measurement.
	Medical Device or Non-Medical Device
	Non-Medical Device

	16th AMDTC

	224.
	Defibrillator analyzers
	Defibrillator analyzers automate the inspection and preventive maintenance (IPM) testing of defibrillators. They need to be able to test at least four basic defibrillator performance characteristics: discharge energy, synchronized-mode operation, automated external defibrillation, and ECG monitoring.
	Medical Device or Non-Medical Device
	Non-Medical Device

	225.
	Eyeglasses
	Eyeglasses with corrective power consists of
- frame (non-medical device)
- lens with power (already registered as medical device)
These two components are assembled and supplied together as a finished product.
	Medical Device or Non-Medical Device
	Medical Device

	226.
	Antiseptic swab with lidocaine
	A pre-moistened swab that combines alcohol and lidocaine for topical application.
	Medical Device or Non-Medical Device
	Non-Medical Device 

	227.
	Distance Vision Acuity Chart
	It is designed for testing distance visual acuity in children and adults who have difficulty with visual acuity charts (within 10 feet / 3 meters). It is intended to assist vision assessment
	Medical Device or Non-Medical Device
	Medical Device 

	228.
	Vision Screening chart
	Vision Screening chart is for screening the vision of individuals beginning at age 3 through adulthood. Spiral-bound, multisided, distance visual acuity eye charts that can be hung on a wall or held in the hands.
	Medical Device or Non-Medical Device
	Medical Device 

	229.
	Vision Colour Test Chart / Book / Card
	An optical instrument color test which is to conduct screening or examination of color deficiency for infants/adults.
	Medical Device or Non-Medical Device
	Medical Device

	230.
	Low Vision and Low Contrast Chart / Card / Book
	To measure or examine or assist visual acuity, assist vision assessment by determining how well eyes function in low light situations and/or the ability to distinguish objects from similarly colored or shaded backgrounds
	Medical Device or Non-Medical Device
	Medical Device

	231.
	Orthosis Support
	An external support which is intended for maximal joint protection, protection for ligament and providing/enhancing joint stabilization and provide adequate range of motion during activity (e.g. finger support, back support, elbow support)
	Medical Device or Non-Medical Device
	Medical Device 

	232.
	Waste bag
	Waste bag refers to a disposable or reusable container, typically made of plastic or other durable materials, designed for the safe collection, containment, and disposal of various types of waste, including general, medical, hazardous, or biological waste. It is commonly used in healthcare facilities, laboratories, households, and industrial settings to promote hygiene, prevent contamination, and support proper waste management practices.
	Medical Device or Non-Medical Device 
	Non-Medical Device




PAGE  
40

